## EB010-EB210 Graph
EB010/EB210 Energy Hand Device Yield
[5 empty] | 0 | Scrap Breakdown (Last 3 S/O)
[6 empty] | Description | Count
[6 empty] | Stuck Blade | 188
[6 empty] | Rough Actuation of Jaw | 179
[6 empty] | Incorrect Insulation Orientation | 65
[6 empty] | Insulation Damage | 59
[6 empty] | 0.003 Jaw Gap Fail | 28
[6 empty] | 0.006 Jaw Gap Fail | 25
[6 empty] | Continuity Fail | 16
[6 empty] | Front Stops Not Touching | 7
[6 empty] | Jaw Short | 5
[6 empty] | Middle Stops Not Touching | 3
[6 empty] | Blade Does Not Pass Front Stop | 1
[6 empty] | Fuse Switch | 1
[6 empty] | Low Jaw Force | 1
[6 empty] | Misassembled | 0
[6 empty] | Damaged Harness | 0
[6 empty] | High Jaw Force | 0
| Shop Order's Yield |  |  |  |  | Rough Knob Actuation | 0
Shop Order | Build QTY | Yield | Yield % | Date |  | Return Spring Pop Out | 0
1484983 | 2187 | 1821 | 83.26% | 2023-04-17 |  | Blade Lever Stop Pop Out | 0
1486323 | 2225 | 1833 | 82.38% | 2023-05-05 |  | Blade Activation | 0
1486218 | 2145 | 1858 | 86.62% | 2023-05-15 |  | Both Stops Not Touching | 0
1486755 | 2297 | 1890 | 82.28% | 2023-05-26 |  | Open Handles | 0
1486760 | 2074 | 1899 | 91.56% | 2023-06-08 |  | Knob Collars Incorrectly Pressed | 0
1488029 | 2092 | 1895 | 90.58% | 2023-06-20 |  | Rough Trigger | 0
Average Yield: |  |  | 86.12% |  |  | Squeaky Blade | 0

## EB010-EB210
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1481234 | EB210 | 1920 | 2142 | 1850 | 86.37% | 2023-04-05 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 16 |  | 13 |  | 5 | [6 empty] | 11 | 45 | 2.101% | 2142 | Stuck Blade | 45
[7 empty] | 44 |  | 1 |  | 1 | [6 empty] | 3 | 49 | 2.288% | 2142 | Rough Actuation of Jaw | 49
[7 empty] | 22 |  | 3 | [8 empty] | 5 | 30 | 1.401% | 2142 | 0.003 Jaw Gap Fail | 30
[7 empty] | 11 |  | 1 | [8 empty] | 4 | 16 | 0.747% | 2142 | 0.006 Jaw Gap Fail | 16
[7 empty] | 4 |  | 1 | [9 empty] | 5 | 0.2334% | 2142 | Front Stops Not Touching | 5
[7 empty] | 33 | [11 empty] | 33 | 1.541% | 2142 | Uneven Jaw | 33
[11 empty] | 11 | [7 empty] | 11 | 0.5135% | 2142 | Damaged Harness | 11
[7 empty] | 3 | [11 empty] | 3 | 0.1401% | 2142 | Blade Does Not Pass Front Stop | 3
[7 empty] | 2 | [11 empty] | 2 | 0.09337% | 2142 | Jaw Short | 2
[7 empty] | 9 |  | 1 | [8 empty] | 11 | 21 | 0.9804% | 2142 | Insulation Damage | 21
[7 empty] | 14 |  | 1 | [9 empty] | 15 | 0.7003% | 2142 | Incorrect Insulation Orientation | 15
[19 empty] | 0 | 0% | 2142 | Missing Harness | 0 | Under Issue
[19 empty] | 0 | 0% | 2142 | Missing Component | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1401% | 2142 | Incorrect/Missing Weld | 3
[7 empty] | 3 |  | 1 | [9 empty] | 4 | 0.1867% | 2142 | Knob Collar Press Incorrectly | 4
[8 empty] | 13 |  | 2 | 2 | [7 empty] | 2 | 0.09337% | 2142 | Rough Knob Actuation | 2
[19 empty] | 0 | 0% | 2142 | Squeaky Blade | 0
[8 empty] | 4 | 1 | 2 | 2 | [6 empty] | 1 | 4 | 0.1867% | 2142 | Continuity Fail | 4
[8 empty] | 98 | 5 | 4 | 2 | [6 empty] | 3 | 10 | 0.4669% | 2142 | High Jaw Force | 10
[19 empty] | 0 | 0% | 2142 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2142 | Return Spring Pop Out | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2142 | Fuse Switch | 0
[8 empty] | 1 |  | 1 | [8 empty] | 0 | 0% | 2142 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[8 empty] | 5 | [9 empty] | 9 | 9 | 0.4202% | 2142 | Scratched Logo | 9 | Damaged Rear Blade Pusher: ea
[8 empty] | 5 |  | 3 | [8 empty] | 0 | 0% | 2142 | Misassembled | 0
[10 empty] | 1 | [8 empty] | 0 | 0% | 2142 | Missing Component | 0
[19 empty] | 0 | 0% | 2142 | Failure To Unlock Trigger | 0
[8 empty] | 5 | [10 empty] | 0 | 0% | 2142 | Cosmetic Defect | 0
[22 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2142 | Incorrect Hub Press | 0 | 7ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2142 | Incorrect Jaw Crimp | 0 | 25ea Incorrect Blade Stop Crimp (No Jaw)
[19 empty] | 0 | 0% | 2142 | Exposed Wire | 0 | 2ea Damaged Cover Tube
[19 empty] | 0 | 0% | Build QTY | Damaged Nut Thread | 0 | 3ea Damaged Pull Tube
[19 empty] | 0 | 0% | 2142 | Incorrect Blade Stop Crimp | 0 | 1ea Damaged Upper Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.09337% | 2142 | Stuck Blade | 2
[19 empty] | 0 | 0% | 2142 | Damaged Insulation | 0
[7 empty] | 21 | [11 empty] | 21 | 0.9804% | 2142 | Damaged Blade | 21
[7 empty] | 7 | [11 empty] | 7 | 0.3268% | 2142 | Incorrect/Missing Weld | 7
[19 empty] | 0 | 0% | 2142 | Blade Pusher Coming Out | 0
[6 empty] | Total | 194 | 133 | 28 | 13 | 23 | 0 | 0 | 0 | 0 | 0 | 0 | 47 | 292 | 13.63% | 2142
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486999 | EB210 | 1920 | 2040 | 1811 | 88.77% | 2023-04-11 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 13 |  | 12 | [8 empty] | 8 | 33 | 1.618% | 2040 | Stuck Blade | 33
[7 empty] | 40 |  | 1 | [8 empty] | 14 | 55 | 2.696% | 2040 | Rough Actuation of Jaw | 55
[7 empty] | 13 |  | 1 | [9 empty] | 14 | 0.6863% | 2040 | 0.003 Jaw Gap Fail | 14
[7 empty] | 8 |  | 1 | [8 empty] | 2 | 11 | 0.5392% | 2040 | 0.006 Jaw Gap Fail | 11
[7 empty] | 1 | [11 empty] | 1 | 0.04902% | 2040 | Front Stops Not Touching | 1
[7 empty] | 4 | [11 empty] | 4 | 0.1961% | 2040 | Uneven Jaw | 4
[9 empty] | 4 | [9 empty] | 4 | 0.1961% | 2040 | Damaged Harness | 4
[19 empty] | 0 | 0% | 2040 | Blade Does Not Pass Front Stop | 0
[7 empty] | 1 |  | 1 | [9 empty] | 2 | 0.09804% | 2040 | Jaw Short | 2
[7 empty] | 5 |  | 1 | [8 empty] | 9 | 15 | 0.7353% | 2040 | Insulation Damage | 15
[7 empty] | 1 | [11 empty] | 1 | 0.04902% | 2040 | Incorrect Insulation Orientation | 1
[18 empty] | 32 | 32 | 1.569% | 2040 | Engineering Sample | 32
[19 empty] | 0 | 0% | 2040 | Missing Component | 0
[9 empty] | 1 | [9 empty] | 1 | 0.04902% | 2040 | Contamination | 1
[7 empty] | 3 | [11 empty] | 3 | 0.1471% | 2040 | Knob Collar Press Incorrectly | 3
[8 empty] | 5 | 1 | [9 empty] | 1 | 0.04902% | 2040 | Rough Knob Actuation | 1
[19 empty] | 0 | 0% | 2040 | Squeaky Blade | 0
[8 empty] | 6 | 6 | [8 empty] | 3 | 9 | 0.4412% | 2040 | Continuity Fail | 9
[8 empty] | 80 | 8 | [8 empty] | 1 | 9 | 0.4412% | 2040 | High Jaw Force | 9
[8 empty] | 1 | [9 empty] | 1 | 1 | 0.04902% | 2040 | Low Jaw Force | 1
[8 empty] | 13 | [10 empty] | 0 | 0% | 2040 | Return Spring Pop Out | 0
[8 empty] | 1 | [9 empty] | 1 | 1 | 0.04902% | 2040 | Fuse Switch | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2040 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[8 empty] | 1 | [9 empty] | 1 | 1 | 0.04902% | 2040 | Scratched Logo | 1 | Damaged Rear Blade Pusher: 51ea
[8 empty] | 11 | [10 empty] | 0 | 0% | 2040 | Misassembled | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2040 | Missing Component | 0
[19 empty] | 0 | 0% | 2040 | Failure To Unlock Trigger | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2040 | Cosmetic Defect | 0
[22 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2040 | Incorrect Hub Press | 0 | 6ea Incorrect Hub Press
[7 empty] | 3 | [11 empty] | 3 | 0.1471% | 2040 | Incorrect Jaw Crimp | 3 | 7ea Incorrect Blade Stop Crimp (No Jaw)
[7 empty] | 5 | [11 empty] | 5 | 0.2451% | 2040 | Exposed Wire | 5 | 4ea Damaged Pull Tube
[19 empty] | 0 | 0% | Build QTY | Damaged Nut Thread | 0 | 12ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2040 | Incorrect Blade Stop Crimp | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04902% | 2040 | Stuck Blade | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04902% | 2040 | Defective Cover Tube | 1
[7 empty] | 5 | [11 empty] | 5 | 0.2451% | 2040 | Damaged Blade | 5
[19 empty] | 0 | 0% | 2040 | Incorrect/Missing Weld | 0
[7 empty] | 16 | [11 empty] | 16 | 0.7843% | 2040 | Blade Pusher Coming Out | 16
[6 empty] | Total | 120 | 122 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 72 | 229 | 11.23% | 2040
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1484982 | EB210 | 1920 | 2061 | 1824 | 88.5% | 2023-04-13 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 29 |  | 13 | [8 empty] | 15 | 57 | 2.766% | 2061 | Stuck Blade | 57
[7 empty] | 28 |  | 1 | [8 empty] | 6 | 35 | 1.698% | 2061 | Rough Actuation of Jaw | 35
[7 empty] | 21 |  | 1 | [8 empty] | 2 | 24 | 1.164% | 2061 | 0.003 Jaw Gap Fail | 24
[7 empty] | 11 |  | 1 | [8 empty] | 7 | 19 | 0.9219% | 2061 | 0.006 Jaw Gap Fail | 19
[7 empty] | 4 | [11 empty] | 4 | 0.1941% | 2061 | Front Stops Not Touching | 4
[7 empty] | 2 | [11 empty] | 2 | 0.09704% | 2061 | Middle Stops Not Touching | 2
[9 empty] | 3 | [9 empty] | 3 | 0.1456% | 2061 | Damaged Harness | 3
[7 empty] | 2 | [11 empty] | 2 | 0.09704% | 2061 | Blade Does Not Pass Front Stop | 2
[19 empty] | 0 | 0% | 2061 | Jaw Short | 0
[7 empty] | 11 | [10 empty] | 4 | 15 | 0.7278% | 2061 | Insulation Damage | 15
[7 empty] | 5 |  | 1 | [9 empty] | 6 | 0.2911% | 2061 | Incorrect Insulation Orientation | 6
[7 empty] | 7 | [11 empty] | 7 | 0.3396% | 2061 | Incorrect/Missing Weld | 7
[18 empty] | 32 | 32 | 1.553% | 2061 | Engineering Sample | 32
[7 empty] | 2 | [11 empty] | 2 | 0.09704% | 2061 | Uneven Jaw | 2
[7 empty] | 2 | [11 empty] | 2 | 0.09704% | 2061 | Knob Collar Press Incorrectly | 2
[8 empty] | 13 | 4 | [9 empty] | 4 | 0.1941% | 2061 | Rough Knob Actuation | 4
[19 empty] | 0 | 0% | 2061 | Squeaky Blade | 0
[8 empty] | 3 | 4 | [8 empty] | 1 | 5 | 0.2426% | 2061 | Continuity Fail | 5
[8 empty] | 13 | [10 empty] | 0 | 0% | 2061 | High Jaw Force | 0
[18 empty] | 1 | 1 | 0.04852% | 2061 | Low Jaw Force | 1
[19 empty] | 0 | 0% | 2061 | Return Spring Pop Out | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 2061 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2061 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2061 | Scratched Logo | 0 | Damaged Rear Blade Pusher: 101ea
[8 empty] | 6 | [10 empty] | 0 | 0% | 2061 | Misassembled | 0
[19 empty] | 0 | 0% | 2061 | Missing Component | 0
[19 empty] | 0 | 0% | 2061 | Failure To Unlock Trigger | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2061 | Cosmetic Defect | 0
[22 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2061 | Incorrect Hub Press | 0 | 2ea Incorrect Hub Press
[7 empty] | 2 | [11 empty] | 2 | 0.09704% | 2061 | Incorrect Jaw Crimp | 2 | 5ea Incorrect Blade Stop Crimp (No Jaw)
[19 empty] | 0 | 0% | 2061 | Exposed Wire | 0 | 1ea Damaged Pull Tube
[7 empty] | 2 | [11 empty] | 2 | 0.09704% | Build QTY | Damaged Nut Thread | 2
[19 empty] | 0 | 0% | 2061 | Incorrect Blade Stop Crimp | 0
[7 empty] | 2 | [11 empty] | 2 | 0.09704% | 2061 | Stuck Blade | 2
[19 empty] | 0 | 0% | 2061 | Defective Cover Tube | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1456% | 2061 | Damaged Component | 3
[7 empty] | 5 | [11 empty] | 5 | 0.2426% | 2061 | Incorrect/Missing Weld | 5
[7 empty] | 3 | [11 empty] | 3 | 0.1456% | 2061 | Blade Pusher Coming Out | 3
[6 empty] | Total | 141 | 42 | 28 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 68 | 237 | 11.5% | 2061
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1484983 | EB210 | 1920 | 2187 | 1821 | 83.26% | 2023-04-17 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 33 | 32 | 47 |  | 6 | [6 empty] | 6 | 92 | 4.207% | 2187 | Stuck Blade | 92
[7 empty] | 72 | 72 | 3 | [8 empty] | 8 | 83 | 3.795% | 2187 | Rough Actuation of Jaw | 83
[7 empty] | 22 | 22 | 1 | [8 empty] | 2 | 25 | 1.143% | 2187 | 0.003 Jaw Gap Fail | 25
[7 empty] | 3 | 3 | [9 empty] | 3 | 6 | 0.2743% | 2187 | 0.006 Jaw Gap Fail | 6
[7 empty] | 1 | 1 |  |  | 1 | [6 empty] | 1 | 3 | 0.1372% | 2187 | Front Stops Not Touching | 3
[11 empty] | 1 | [7 empty] | 1 | 0.04572% | 2187 | Middle Stops Not Touching | 1
[19 empty] | 0 | 0% | 2187 | Damaged Harness | 0
[7 empty] | 4 | 4 | [10 empty] | 4 | 0.1829% | 2187 | Blade Does Not Pass Front Stop | 4
[7 empty] | 4 | 4 | [9 empty] | 1 | 5 | 0.2286% | 2187 | Jaw Short | 5
[7 empty] | 17 | 17 | [9 empty] | 5 | 22 | 1.006% | 2187 | Insulation Damage | 22
[7 empty] | 1 | 1 | [10 empty] | 1 | 0.04572% | 2187 | Incorrect Insulation Orientation | 1
[7 empty] | 4 | 4 | [10 empty] | 4 | 0.1829% | 2187 | Incorrect/Missing Weld | 4
[18 empty] | 32 | 32 | 1.463% | 2187 | Engineering Sample | 32
[7 empty] | 11 | 11 | [10 empty] | 11 | 0.503% | 2187 | Uneven Jaw | 11
[7 empty] | 18 | 18 | [10 empty] | 18 | 0.823% | 2187 | Knob Collar Press Incorrectly | 18
[8 empty] | 13 | 1 | 1 | 1 | [7 empty] | 2 | 0.09145% | 2187 | Rough Knob Actuation | 2
[19 empty] | 0 | 0% | 2187 | Squeaky Blade | 0
[8 empty] | 8 | 4 | 3 | 1 | [6 empty] | 2 | 7 | 0.3201% | 2187 | Continuity Fail | 7
[8 empty] | 125 | 14 | 7 | 16 | [6 empty] | 1 | 31 | 1.417% | 2187 | High Jaw Force | 31
[19 empty] | 0 | 0% | 2187 | Low Jaw Force | 0
[11 empty] | 1 | [7 empty] | 1 | 0.04572% | 2187 | Return Spring Pop Out | 1
[8 empty] | 5 | [10 empty] | 0 | 0% | 2187 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2187 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[18 empty] | 4 | 4 | 0.1829% | 2187 | Scratched Logo | 4 | Damaged Front Blade Pusher: 89ea
[8 empty] | 18 | 1 |  | 1 | [7 empty] | 2 | 0.09145% | 2187 | Misassembled | 2 | Damaged Insulation : 168ea
[11 empty] | 2 | [7 empty] | 2 | 0.09145% | 2187 | Missing Component | 2
[11 empty] | 2 | [7 empty] | 2 | 0.09145% | 2187 | Contamination | 2
[8 empty] | 2 | [10 empty] | 0 | 0% | 2187 | Cosmetic Defect | 0
[22 empty] | Shafts | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04572% | 2187 | Uneven Jaw | 1 | 16ea Incorrect Hub Press
[7 empty] | 2 | [11 empty] | 2 | 0.09145% | 2187 | Insulation Damage | 2 | 17ea Incorrect Blade Stop Crimp (No Jaw)
[7 empty] | 4 | [11 empty] | 4 | 0.1829% | 2187 | Exposed Wire | 4 | 9ea Damaged Pull Tube
[19 empty] | 0 | 0% | Build QTY | Damaged Nut Thread | 0 | 1ea Damaged Lower jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04572% | 2187 | Incorrect Blade Stop Crimp | 1 | 2ea Damaged Upper Jaw
[7 empty] | 4 | [11 empty] | 4 | 0.1829% | 2187 | Stuck Blade | 4
[19 empty] | 0 | 0% | 2187 | Defective Cover Tube | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1372% | 2187 | Damaged Blade | 3
[19 empty] | 0 | 0% | 2187 | Incorrect/Missing Weld | 0
[7 empty] | 10 | [11 empty] | 10 | 0.4572% | 2187 | Blade Pusher Coming Out | 10
[6 empty] | Total | 215 | 360 | 71 | 11 | 32 | 0 | 0 | 0 | 0 | 0 | 0 | 65 | 383 | 17.51% | 2187
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486323 | EB210 | 1920 | 2225 | 1833 | 82.38% | 2023-05-05 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 45 |  | 27 |  | 3 | [6 empty] | 15 | 90 | 4.045% | 2225 | Stuck Blade | 90
[7 empty] | 48 |  | 4 |  | 1 | [6 empty] | 16 | 69 | 3.101% | 2225 | Rough Actuation of Jaw | 69
[7 empty] | 50 |  | 6 |  | 3 | [6 empty] | 15 | 74 | 3.326% | 2225 | 0.003 Jaw Gap Fail | 74
[7 empty] | 10 |  | 8 | [8 empty] | 2 | 20 | 0.8989% | 2225 | 0.006 Jaw Gap Fail | 20
[7 empty] | 5 | [11 empty] | 5 | 0.2247% | 2225 | Front Stops Not Touching | 5
[19 empty] | 0 | 0% | 2225 | Middle Stops Not Touching | 0
[11 empty] | 5 | [7 empty] | 5 | 0.2247% | 2225 | Damaged Harness | 5
[19 empty] | 0 | 0% | 2225 | Blade Does Not Pass Front Stop | 0
[7 empty] | 8 |  | 1 |  | 1 | [7 empty] | 10 | 0.4494% | 2225 | Jaw Short | 10
[7 empty] | 6 |  | 11 | [8 empty] | 3 | 20 | 0.8989% | 2225 | Insulation Damage | 20
[7 empty] | 39 |  | 2 | [8 empty] | 7 | 48 | 2.157% | 2225 | Incorrect Insulation Orientation | 48
[7 empty] | 2 | [11 empty] | 2 | 0.08989% | 2225 | Incorrect/Missing Weld | 2
[19 empty] | 0 | 0% | 2225 | Engineering Sample | 0
[7 empty] | 4 | [11 empty] | 4 | 0.1798% | 2225 | Uneven Jaw | 4
[7 empty] | 6 |  | 1 | [9 empty] | 7 | 0.3146% | 2225 | Knob Collar Press Incorrectly | 7
[8 empty] | 7 | [10 empty] | 0 | 0% | 2225 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2225 | Squeaky Blade | 0
[8 empty] | 1 | [9 empty] | 5 | 5 | 0.2247% | 2225 | Continuity Fail | 5
[8 empty] | 355 |  |  | 1 | [6 empty] | 8 | 9 | 0.4045% | 2225 | High Jaw Force | 9
[19 empty] | 0 | 0% | 2225 | Low Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2225 | Return Spring Pop Out | 0
[8 empty] | 2 | [9 empty] | 1 | 1 | 0.04494% | 2225 | Fuse Switch | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2225 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 3ea
[8 empty] | 2 | [10 empty] | 0 | 0% | 2225 | Scratched Logo | 0 | Damaged Front Blade Pusher: 39ea
[8 empty] | 7 |  |  | 3 | [7 empty] | 3 | 0.1348% | 2225 | Misassembled | 3 | Damaged Rear Blade Pusher: 14a
[8 empty] | 2 | [10 empty] | 0 | 0% | 2225 | Missing Component | 0 | Damaged Insulation : 110ea
[19 empty] | 0 | 0% | 2225 | Contamination | 0
[8 empty] | 6 | [10 empty] | 0 | 0% | 2225 | Cosmetic Defect | 0
[22 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2225 | Uneven Jaw | 0 | 10ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.04494% | 2225 | Insulation Damage | 1 | 10ea Incorrect Blade Stop Crimp (No Jaw)
[7 empty] | 2 | [11 empty] | 2 | 0.08989% | 2225 | Exposed Wire | 2 | 3ea Damaged Pull Tube
[19 empty] | 0 | 0% | Build QTY | Damaged Nut Thread | 0 | 1ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2225 | Incorrect Blade Stop Crimp | 0 | 6ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2225 | Stuck Blade | 0
[19 empty] | 0 | 0% | 2225 | Defective Cover Tube | 0
[7 empty] | 7 | [11 empty] | 7 | 0.3146% | 2225 | Damaged Blade | 7
[7 empty] | 6 | [11 empty] | 6 | 0.2697% | 2225 | Incorrect/Missing Weld | 6
[7 empty] | 6 | [11 empty] | 6 | 0.2697% | 2225 | Blade Pusher Coming Out | 6
[6 empty] | Total | 245 | 384 | 60 | 0 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 72 | 394 | 17.71% | 2225
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486218 | EB210 | 1920 | 2145 | 1858 | 86.62% | 2023-05-15 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 33 |  | 15 |  | 12 | [6 empty] | 8 | 68 | 3.17% | 2145 | Stuck Blade | 68
[7 empty] | 50 |  | 27 | [8 empty] | 5 | 82 | 3.823% | 2145 | Rough Actuation of Jaw | 82
[7 empty] | 10 |  |  |  | 2 | [6 empty] | 3 | 15 | 0.6993% | 2145 | 0.003 Jaw Gap Fail | 15
[7 empty] | 5 |  |  |  | 1 | [6 empty] | 3 | 9 | 0.4196% | 2145 | 0.006 Jaw Gap Fail | 9
[7 empty] | 4 | [10 empty] | 1 | 5 | 0.2331% | 2145 | Front Stops Not Touching | 5
[7 empty] | 1 | [11 empty] | 1 | 0.04662% | 2145 | Middle Stops Not Touching | 1
[7 empty] | 1 |  |  |  | 3 | [7 empty] | 4 | 0.1865% | 2145 | Damaged Harness | 4
[7 empty] | 2 | [11 empty] | 2 | 0.09324% | 2145 | Blade Does Not Pass Front Stop | 2
[7 empty] | 2 | [10 empty] | 8 | 10 | 0.4662% | 2145 | Jaw Short | 10
[7 empty] | 11 | [11 empty] | 11 | 0.5128% | 2145 | Insulation Damage | 11
[7 empty] | 21 |  | 13 |  | 1 | [6 empty] | 5 | 40 | 1.865% | 2145 | Incorrect Insulation Orientation | 40
[7 empty] | 1 | [11 empty] | 1 | 0.04662% | 2145 | Incorrect/Missing Weld | 1
[7 empty] | 2 | [11 empty] | 2 | 0.09324% | 2145 | Missing Jaw Crimp | 2
[7 empty] | 7 |  | 2 | [9 empty] | 9 | 0.4196% | 2145 | Uneven Jaw | 9
[7 empty] | 3 | [11 empty] | 3 | 0.1399% | 2145 | Knob Collar Press Incorrectly | 3
[19 empty] | 0 | 0% | 2145 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2145 | Squeaky Blade | 0
[8 empty] | 3 | 1 | [8 empty] | 2 | 3 | 0.1399% | 2145 | Continuity Fail | 3
[8 empty] | 1 |  | 1 | [8 empty] | 0 | 0% | 2145 | High Jaw Force | 0
[19 empty] | 0 | 0% | 2145 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2145 | Return Spring Pop Out | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2145 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2145 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[18 empty] | 5 | 5 | 0.2331% | 2145 | Scratched Logo | 5 | Damaged Front Blade Pusher: 72ea
[8 empty] | 10 | 2 | 5 | 3 | [7 empty] | 5 | 0.2331% | 2145 | Misassembled | 5 | Damaged Insulation : 204ea
[19 empty] | 0 | 0% | 2145 | Missing Component | 0
[19 empty] | 0 | 0% | 2145 | Contamination | 0
[7 empty] | 3 | 2 | [10 empty] | 3 | 0.1399% | 2145 | Cosmetic Defect | 3
[22 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2145 | Uneven Jaw | 0 | 2ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.04662% | 2145 | Insulation Damage | 1 | 4ea Incorrect Blade Stop Crimp (No Jaw)
[7 empty] | 2 | [11 empty] | 2 | 0.09324% | 2145 | Exposed Wire | 2 | 8ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04662% | Build QTY | Damaged Component | 1 | 2ea Contamination Cover Tube
[19 empty] | 0 | 0% | 2145 | Incorrect Blade Stop Crimp | 0 | 10ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04662% | 2145 | Stuck Blade | 1
[19 empty] | 0 | 0% | 2145 | Defective Cover Tube | 0
[7 empty] | 2 | [11 empty] | 2 | 0.09324% | 2145 | Damaged Blade | 2
[7 empty] | 1 | [11 empty] | 1 | 0.04662% | 2145 | Incorrect/Missing Weld | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04662% | 2145 | Blade Pusher Coming Out | 1
[6 empty] | Total | 165 | 17 | 60 | 6 | 22 | 0 | 0 | 0 | 0 | 0 | 0 | 40 | 287 | 13.38% | 2145
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486755 | EB210 | 1920 | 2297 | 1890 | 82.28% | 2023-05-26 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 48 |  | 8 | [8 empty] | 4 | 60 | 2.612% | 2297 | Stuck Blade | 60
[7 empty] | 93 |  | 2 | [8 empty] | 14 | 109 | 4.745% | 2297 | Rough Actuation of Jaw | 109
[7 empty] | 12 |  | 4 | [8 empty] | 3 | 19 | 0.8272% | 2297 | 0.003 Jaw Gap Fail | 19
[7 empty] | 2 |  | 1 | [8 empty] | 2 | 5 | 0.2177% | 2297 | 0.006 Jaw Gap Fail | 5
[19 empty] | 0 | 0% | 2297 | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 2297 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2297 | Damaged Harness | 0
[19 empty] | 0 | 0% | 2297 | Blade Does Not Pass Front Stop | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04354% | 2297 | Jaw Short | 1
[7 empty] | 22 | [10 empty] | 7 | 29 | 1.263% | 2297 | Insulation Damage | 29
[7 empty] | 35 |  | 4 | [9 empty] | 39 | 1.698% | 2297 | Incorrect Insulation Orientation | 39
[7 empty] | 5 | [11 empty] | 5 | 0.2177% | 2297 | Incorrect/Missing Weld | 5
[7 empty] | 26 | [11 empty] | 26 | 1.132% | 2297 | Engineer Sample | 26 | No Harness
[7 empty] | 7 |  | 7 | [9 empty] | 14 | 0.6095% | 2297 | Uneven Jaw | 14
[7 empty] | 4 |  | 2 | [9 empty] | 6 | 0.2612% | 2297 | Knob Collar Press Incorrectly | 6
[10 empty] | 1 | [8 empty] | 0 | 0% | 2297 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2297 | Squeaky Blade | 0
[8 empty] | 4 | 1 | 2 | [7 empty] | 1 | 2 | 0.08707% | 2297 | Continuity Fail | 2
[8 empty] | 2 | [10 empty] | 0 | 0% | 2297 | High Jaw Force | 0
[19 empty] | 0 | 0% | 2297 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2297 | Return Spring Pop Out | 0
[19 empty] | 0 | 0% | 2297 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2297 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[18 empty] | 2 | 2 | 0.08707% | 2297 | Scratched Logo | 2 | Damaged Front Blade Pusher: 85ea
[8 empty] | 14 |  | 2 | [8 empty] | 0 | 0% | 2297 | Misassembled | 0 | Damaged Insulation : 117ea
[19 empty] | 0 | 0% | 2297 | Missing Component | 0
[19 empty] | 0 | 0% | 2297 | Contamination | 0
[8 empty] | 12 | [10 empty] | 0 | 0% | 2297 | Cosmetic Defect | 0
[22 empty] | Shafts | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04354% | 2297 | Contamination | 1 | 3ea Incorrect Hub Press
[7 empty] | 4 | [11 empty] | 4 | 0.1741% | 2297 | Insulation Damage | 4 | 5ea Incorrect Blade Stop Crimp (No Jaw)
[7 empty] | 2 | [11 empty] | 2 | 0.08707% | 2297 | Exposed Wire | 2 | 8ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04354% | Build QTY | Damaged Component | 1 | 9ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04354% | 2297 | Incorrect Blade Stop Crimp | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04354% | 2297 | Blade Does Not Pass Front Stop | 1
[19 empty] | 0 | 0% | 2297 | Defective Cover Tube | 0
[7 empty] | 10 | [11 empty] | 10 | 0.4354% | 2297 | Damaged Blade | 10
[7 empty] | 73 | [11 empty] | 73 | 3.178% | 2297 | Incorrect/Missing Weld | 73
[19 empty] | 0 | 0% | 2297 | Blade Pusher Coming Out | 0
[6 empty] | Total | 348 | 32 | 29 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 33 | 410 | 17.85% | 2297
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486760 | EB210 | 1920 | 2074 | 1899 | 91.56% | 2023-06-08 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 45 |  | 5 | [8 empty] | 7 | 57 | 2.748% | 2074 | Stuck Blade | 57
[7 empty] | 21 |  | 2 | [8 empty] | 2 | 25 | 1.205% | 2074 | Rough Actuation of Jaw | 25
[7 empty] | 5 | [10 empty] | 1 | 6 | 0.2893% | 2074 | 0.003 Jaw Gap Fail | 6
[7 empty] | 6 |  | 1 | [9 empty] | 7 | 0.3375% | 2074 | 0.006 Jaw Gap Fail | 7
[7 empty] | 3 | [11 empty] | 3 | 0.1446% | 2074 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 2074 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2074 | Damaged Harness | 0
[19 empty] | 0 | 0% | 2074 | Blade Does Not Pass Front Stop | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04822% | 2074 | Jaw Short | 1
[7 empty] | 22 | [10 empty] | 1 | 23 | 1.109% | 2074 | Insulation Damage | 23
[7 empty] | 11 | [11 empty] | 11 | 0.5304% | 2074 | Incorrect Insulation Orientation | 11
[19 empty] | 0 | 0% | 2074 | Incorrect/Missing Weld | 0
[11 empty] | 3 | [7 empty] | 3 | 0.1446% | 2074 | Dameged Harness | 3
[7 empty] | 10 | [11 empty] | 10 | 0.4822% | 2074 | Uneven Jaw | 10
[7 empty] | 2 | [11 empty] | 2 | 0.09643% | 2074 | Knob Collar Press Incorrectly | 2
[19 empty] | 0 | 0% | 2074 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2074 | Squeaky Blade | 0
[8 empty] | 7 | 4 | 1 | [7 empty] | 3 | 7 | 0.3375% | 2074 | Continuity Fail | 7
[8 empty] | 2 | [10 empty] | 0 | 0% | 2074 | High Jaw Force | 0
[19 empty] | 0 | 0% | 2074 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2074 | Return Spring Pop Out | 0
[8 empty] | 3 |  | 1 | 1 | [7 empty] | 1 | 0.04822% | 2074 | Fuse Switch | 1
[19 empty] | 0 | 0% | 2074 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[18 empty] | 2 | 2 | 0.09643% | 2074 | Scratched Logo | 2 | Damaged Front Blade Pusher: 81ea
[8 empty] | 8 | [10 empty] | 0 | 0% | 2074 | Misassembled | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2074 | Missing Component | 0
[19 empty] | 0 | 0% | 2074 | Contamination | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2074 | Cosmetic Defect | 0
[22 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2074 | Contamination | 0 | 7ea Incorrect Hub Press
[7 empty] | 3 | [11 empty] | 3 | 0.1446% | 2074 | Incorrect Jaw Crimp | 3 | 5ea Incorrect Blade Stop Crimp (No Jaw)
[19 empty] | 0 | 0% | 2074 | Exposed Wire | 0 | 14ea Damaged Pull Tube
[19 empty] | 0 | 0% | Build QTY | Damaged Component | 0 | 1ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2074 | Incorrect Blade Stop Crimp | 0 | 11ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2074 | Blade Does Not Pass Front Stop | 0
[7 empty] | 4 | [11 empty] | 4 | 0.1929% | 2074 | Stuck Blade | 4
[7 empty] | 4 | [11 empty] | 4 | 0.1929% | 2074 | Damaged Blade | 4
[7 empty] | 5 | [11 empty] | 5 | 0.2411% | 2074 | Incorrect/Missing Weld | 5
[19 empty] | 0 | 0% | 2074 | Blade Pusher Coming Out | 0
[6 empty] | Total | 142 | 22 | 12 | 2 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 174 | 8.39% | 2074
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1488029 | EB210 | 1920 | 2092 | 1895 | 90.58% | 2023-06-20 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 64 | [10 empty] | 3 | 67 | 3.203% | 2092 | Stuck Blade | 67
[7 empty] | 45 | [11 empty] | 45 | 2.151% | 2092 | Rough Actuation of Jaw | 45
[7 empty] | 2 | [10 empty] | 1 | 3 | 0.1434% | 2092 | 0.003 Jaw Gap Fail | 3
[7 empty] | 11 | [10 empty] | 2 | 13 | 0.6214% | 2092 | 0.006 Jaw Gap Fail | 13
[7 empty] | 4 | [11 empty] | 4 | 0.1912% | 2092 | Front Stops Not Touching | 4
[7 empty] | 3 | [11 empty] | 3 | 0.1434% | 2092 | Middle Stops Not Touching | 3
[19 empty] | 0 | 0% | 2092 | Missing Harness | 0
[19 empty] | 0 | 0% | 2092 | Blade Does Not Pass Front Stop | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1434% | 2092 | Jaw Short | 3
[7 empty] | 3 | [11 empty] | 3 | 0.1434% | 2092 | Insulation Damage | 3
[7 empty] | 11 | [10 empty] | 4 | 15 | 0.717% | 2092 | Incorrect Insulation Orientation | 15
[7 empty] | 2 | [11 empty] | 2 | 0.0956% | 2092 | Incorrect/Missing Weld | 2
[9 empty] | 3 | [9 empty] | 3 | 0.1434% | 2092 | Dameged Harness | 3
[7 empty] | 3 | [11 empty] | 3 | 0.1434% | 2092 | Uneven Jaw | 3
[7 empty] | 4 | [11 empty] | 4 | 0.1912% | 2092 | Knob Collar Press Incorrectly | 4
[8 empty] | 2 | [10 empty] | 0 | 0% | 2092 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2092 | Squeaky Blade | 0
[8 empty] | 4 | [9 empty] | 7 | 7 | 0.3346% | 2092 | Continuity Fail | 7
[19 empty] | 0 | 0% | 2092 | High Jaw Force | 0
[18 empty] | 1 | 1 | 0.0478% | 2092 | Low Jaw Force | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2092 | Return Spring Pop Out | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2092 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2092 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2092 | Scratched Logo | 0 | Damaged Front Blade Pusher: 77ea
[8 empty] | 8 | [10 empty] | 0 | 0% | 2092 | Misassembled | 0 | Insulation Damage: 132ea
[8 empty] | 1 | [10 empty] | 0 | 0% | 2092 | Missing Component | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2092 | Wire Pinch | 0
[8 empty] | 8 | 4 | [9 empty] | 4 | 0.1912% | 2092 | Cosmetic Defect | 4
[22 empty] | Shafts | 0
[7 empty] | 1 | [11 empty] | 1 | 0.0478% | 2092 | Contamination | 1 | 5ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.0478% | 2092 | Incorrect Jaw Crimp | 1 | 2ea Incorrect Blade Stop Crimp (No Jaw)
[7 empty] | 1 | [11 empty] | 1 | 0.0478% | 2092 | Exposed Wire | 1 | 9ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.0478% | Build QTY | Damaged Component | 1 | 11ea Contamination Lower Jaw
[19 empty] | 0 | 0% | 2092 | Incorrect Blade Stop Crimp | 0 | 6ea Exposed Wire Upper Jaw
[19 empty] | 0 | 0% | 2092 | Blade Does Not Pass Front Stop | 0
[19 empty] | 0 | 0% | 2092 | Stuck Blade | 0
[7 empty] | 6 | [11 empty] | 6 | 0.2868% | 2092 | Damaged Blade | 6
[19 empty] | 0 | 0% | 2092 | Incorrect/Missing Weld | 0
[7 empty] | 7 | [11 empty] | 7 | 0.3346% | 2092 | Blade Pusher Coming Out | 7
[6 empty] | Total | 172 | 27 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 197 | 9.417% | 2092

## EB011-EB211 Graphs
EB011/EB211 Energy Hand Device Yield
Scrap Breakdown (Last S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 51
[5 empty] | Rough Actuation of Jaw | 44
[5 empty] | Insulation Damage | 22
[5 empty] | Incorrect Insulation Orientation | 12
[5 empty] | 0.006 Jaw Gap Fail | 8
[5 empty] | Continuity Fail | 6
[5 empty] | 0.003 Jaw Gap Fail | 5
[5 empty] | Front Stops Not Touching | 5
[5 empty] | Jaw Short | 2
[5 empty] | Middle Stops Not Touching | 1
[5 empty] | High Jaw Force | 0
[5 empty] | Low Jaw Force | 0
[5 empty] | Fuse Switch | 0
[5 empty] | Blade Activation | 0
[5 empty] | Misassembled | 0
[5 empty] | Blade Does Not Pass Front Stop | 0
Shop Order's Yield |  |  |  |  | Misaligned Pins | 0
Shop Order | Build QTY | Yield | Yield % | Date | Gamma Audit | 0
1486756 | 312 | 279 | 89.42% | 2023-04-18 | Squeaky Blade | 0
1486324 | 303 | 277 | 91.42% | 2023-05-08 | Both Stops Not Touching | 0
1490522 | 331 | 279 | 84.29% | 2023-05-20 | Rough Knob Actuation | 0
1492703 | 1192 | 1090 | 91.44% | 2023-06-15
| #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A
Average Yield: |  |  | 89.14%

## EB011-EB211
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1486756 | EB211 | 288 | 312 | 279 | 89.42% | 2023-04-18 | [16 empty] | Top Level Non-Conformance | Total | **Production Shop Order**
[7 empty] | 7 | [11 empty] | 1 | 8 | 2.564% | 312 | Stuck Blade | 8
[7 empty] | 6 | [12 empty] | 6 | 1.923% | 312 | Rough Actuation of Jaw | 6
[7 empty] | 4 | [12 empty] | 4 | 1.282% | 312 | 0.003 Jaw Gap Fail | 4
[19 empty] | 1 | 1 | 0.3205% | 312 | 0.006 Jaw Gap Fail | 1
[7 empty] | 2 | [12 empty] | 2 | 0.641% | 312 | Front Stops Not Touching | 2
[20 empty] | 0 | 0% | 312 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 312 | Both Stops Not Touching | 0
[20 empty] | 0 | 0% | 312 | Blade Does Not Pass Front Stop | 0
[7 empty] | 1 | [12 empty] | 1 | 0.3205% | 312 | Jaw Short | 1
[7 empty] | 1 | [12 empty] | 1 | 0.3205% | 312 | Insulation Damage | 1
[7 empty] | 1 | [12 empty] | 1 | 0.3205% | 312 | Incorrect Insulation Orientation | 1
[9 empty] | 1 | [10 empty] | 1 | 0.3205% | 312 | Damaged Harness | 1
[20 empty] | 0 | 0% |  | Knob Collar Press Incorrectly
[20 empty] | 0 | 0% | 312 | Damaged Component
[20 empty] | 0 | 0% | 312 | Missing Harness | 0
[8 empty] | 2 | 2 | [10 empty] | 2 | 0.641% | 312 | Rough Knob Actuation
[20 empty] | 0 | 0% | 312 | Squeaky Blade | 0 | BP Cap Separation:
[8 empty] | 3 | 1 | [10 empty] | 1 | 0.3205% | 312 | Continuity Fail | 1 | Damaged Blade Pusher:
[8 empty] | 9 | 3 | [10 empty] | 3 | 0.9615% | 312 | High Jaw Force | 3
[20 empty] | 0 | 0% | 312 | Low Jaw Force | 0
[20 empty] | 0 | 0% | 312 | Return Spring Pop Out | 0
[20 empty] | 0 | 0% | 312 | Fuse Switch | 0
[20 empty] | 0 | 0% | 312 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 312 | Scratched Logo | 0
[8 empty] | 1 | [11 empty] | 0 | 0% | 312 | Misassembled | 0
[20 empty] | 0 | 0% | 312 | Trigger Actuation | 0
[20 empty] | 0 | 0% | 312 | Cosmetic Defect | 0
[20 empty] | 0 | 0% | 312 | Missing Component | 0
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 312 | Cosmetic Defect | 0
[20 empty] | 0 | 0% | 312 | Incorrect Jaw Crimp | 0 | 1ea Incorrect Hub Press
[20 empty] | 0 | 0% | 312 | Missing Component | 0 | 1ea Damaged cover Tube
[7 empty] | 2 | [12 empty] | 2 | 0.641% | 312 | Incorrect Weld | 2 | 1ea Damaged Pull Tube
[20 empty] | 0 | 0% | 312 | Damaged Component | 0
[20 empty] | 0 | 0% | 312 | Contamination | 0
[20 empty] | 0 | 0% | 312 | Exposed Wire | 0
[6 empty] | Total | 24 | 15 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 33 | 10.58% | 312 |  | 33
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1486324 | EB211 | 288 | 303 | 277 | 91.42% | 2023-05-08 | [16 empty] | Top Level Non-Conformance | Total | **Production Shop Order**
[7 empty] | 4 |  | 1 | [9 empty] | 2 | 7 | 2.31% | 303 | Stuck Blade | 7
[7 empty] | 1 | [11 empty] | 2 | 3 | 0.9901% | 303 | Rough Actuation of Jaw | 3
[7 empty] | 1 | [12 empty] | 1 | 0.33% | 303 | 0.003 Jaw Gap Fail | 1
[7 empty] | 2 | [11 empty] | 1 | 3 | 0.9901% | 303 | 0.006 Jaw Gap Fail | 3
[7 empty] | 3 | [12 empty] | 3 | 0.9901% | 303 | Front Stops Not Touching | 3
[20 empty] | 0 | 0% | 303 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 303 | Both Stops Not Touching | 0
[20 empty] | 0 | 0% | 303 | Blade Does Not Pass Front Stop | 0
[20 empty] | 0 | 0% | 303 | Jaw Short | 0
[18 empty] | 1 |  | 0 | 0% | 303 | Insulation Damage | 0
[7 empty] | 1 | [12 empty] | 1 | 0.33% | 303 | Incorrect Insulation Orientation | 1
[20 empty] | 0 | 0% | 303 | Damaged Harness | 0
[20 empty] | 0 | 0% |  | Knob Collar Press Incorrectly
[20 empty] | 0 | 0% | 303 | Damaged Component
[9 empty] | 3 | [10 empty] | 3 | 0.9901% | 303 | Missing Harness | 3
[20 empty] | 0 | 0% | 303 | Rough Knob Actuation
[20 empty] | 0 | 0% | 303 | Squeaky Blade | 0 | BP Cap Separation:
[20 empty] | 0 | 0% | 303 | Continuity Fail | 0 | Damaged Blade Pusher:
[20 empty] | 0 | 0% | 303 | High Jaw Force | 0
[20 empty] | 0 | 0% | 303 | Low Jaw Force | 0
[20 empty] | 0 | 0% | 303 | Return Spring Pop Out | 0
[20 empty] | 0 | 0% | 303 | Fuse Switch | 0
[20 empty] | 0 | 0% | 303 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 303 | Scratched Logo | 0
[20 empty] | 0 | 0% | 303 | Misassembled | 0
[20 empty] | 0 | 0% | 303 | Trigger Actuation | 0
[8 empty] | 2 | 1 | [10 empty] | 1 | 0.33% | 303 | Cosmetic Defect | 1
[20 empty] | 0 | 0% | 303 | Missing Component | 0
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 303 | Cosmetic Defect | 0 | 2ea Incorrect Blade Stop Crimp (No Jaw)
[20 empty] | 0 | 0% | 303 | Incorrect Jaw Crimp | 0 | 1ea Damaged Upper Jaw
[20 empty] | 0 | 0% | 303 | Missing Component | 0
[7 empty] | 1 | [12 empty] | 1 | 0.33% | 303 | Incorrect Weld | 1
[7 empty] | 1 | [12 empty] | 1 | 0.33% | 303 | Damaged Component | 1
[7 empty] | 1 | [12 empty] | 1 | 0.33% | 303 | Stuck Blade | 1
[20 empty] | 0 | 0% | 303 | Exposed Wire | 0
[6 empty] | Total | 15 | 2 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 5 | 25 | 8.251% | 303 |  | 25
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1490522 | EB211 | 288 | 331 | 279 | 84.29% | 2023-05-20 | [16 empty] | Top Level Non-Conformance | Total | **Production Shop Order**
[7 empty] | 25 |  | 2 | [10 empty] | 27 | 8.157% | 331 | Stuck Blade | 27
[7 empty] | 8 |  | 1 | [10 empty] | 9 | 2.719% | 331 | Rough Actuation of Jaw | 9
[20 empty] | 0 | 0% | 331 | 0.003 Jaw Gap Fail | 0
[19 empty] | 1 | 1 | 0.3021% | 331 | 0.006 Jaw Gap Fail | 1
[20 empty] | 0 | 0% | 331 | Front Stops Not Touching | 0
[7 empty] | 1 | [12 empty] | 1 | 0.3021% | 331 | Middle Stops Not Touching | 1
[20 empty] | 0 | 0% | 331 | Both Stops Not Touching | 0
[20 empty] | 0 | 0% | 331 | Blade Does Not Pass Front Stop | 0
[20 empty] | 0 | 0% | 331 | Jaw Short | 0
[7 empty] | 3 | [12 empty] | 3 | 0.9063% | 331 | Insulation Damage | 3
[7 empty] | 1 | [12 empty] | 1 | 0.3021% | 331 | Incorrect Insulation Orientation | 1
[9 empty] | 2 | [10 empty] | 2 | 0.6042% | 331 | Damaged Harness | 2
[7 empty] | 1 | [12 empty] | 1 | 0.3021% |  | Knob Collar Press Incorrectly
[7 empty] | 1 | [12 empty] | 1 | 0.3021% | 331 | Uneven Jaw
[20 empty] | 0 | 0% | 331 | Missing Harness | 0
[20 empty] | 0 | 0% | 331 | Rough Knob Actuation
[20 empty] | 0 | 0% | 331 | Squeaky Blade | 0 | BP Cap Separation:
[19 empty] | 3 | 3 | 0.9063% | 331 | Continuity Fail | 3 | Damaged Blade Pusher: 30ea
[20 empty] | 0 | 0% | 331 | High Jaw Force | 0
[8 empty] | 1 | [11 empty] | 0 | 0% | 331 | Low Jaw Force | 0
[20 empty] | 0 | 0% | 331 | Return Spring Pop Out | 0
[20 empty] | 0 | 0% | 331 | Fuse Switch | 0
[20 empty] | 0 | 0% | 331 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 331 | Scratched Logo | 0
[8 empty] | 4 | [11 empty] | 0 | 0% | 331 | Misassembled | 0
[20 empty] | 0 | 0% | 331 | Trigger Actuation | 0
[20 empty] | 0 | 0% | 331 | Cosmetic Defect | 0
[8 empty] | 1 | [11 empty] | 0 | 0% | 331 | Missing Component | 0
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 331 | Cosmetic Defect | 0
[20 empty] | 0 | 0% | 331 | Incorrect Jaw Crimp | 0 | 1ea Damaged Upper Jaw
[7 empty] | 3 | [12 empty] | 3 | 0.9063% | 331 | Insulation Damage | 3
[20 empty] | 0 | 0% | 331 | Incorrect Weld | 0
[20 empty] | 0 | 0% | 331 | Damaged Component | 0
[20 empty] | 0 | 0% | 331 | Stuck Blade | 0
[20 empty] | 0 | 0% | 331 | Exposed Wire | 0
[6 empty] | Total | 43 | 6 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 52 | 15.71% | 331 |  | 52
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1492703 | EB211 | 1152 | 1192 | 1090 | 91.44% | 2023-06-15 | [16 empty] | Top Level Non-Conformance | Total | **Production Shop Order**
[7 empty] | 7 |  | 3 | [9 empty] | 4 | 14 | 1.174% | 1192 | Stuck Blade | 14
[7 empty] | 7 |  | 19 | [9 empty] | 6 | 32 | 2.685% | 1192 | Rough Actuation of Jaw | 32
[7 empty] | 2 |  | 1 | [9 empty] | 1 | 4 | 0.3356% | 1192 | 0.003 Jaw Gap Fail | 4
[19 empty] | 4 | 4 | 0.3356% | 1192 | 0.006 Jaw Gap Fail | 4
[7 empty] | 1 |  | 1 | [10 empty] | 2 | 0.1678% | 1192 | Front Stops Not Touching | 2
[20 empty] | 0 | 0% | 1192 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 1192 | Both Stops Not Touching | 0
[20 empty] | 0 | 0% | 1192 | Blade Does Not Pass Front Stop | 0
[19 empty] | 2 | 2 | 0.1678% | 1192 | Jaw Short | 2
[7 empty] | 2 |  | 6 | [10 empty] | 8 | 0.6711% | 1192 | Insulation Damage | 8
[7 empty] | 5 |  | 4 | [9 empty] | 1 | 10 | 0.8389% | 1192 | Incorrect Insulation Orientation | 10
[9 empty] | 1 | [10 empty] | 1 | 0.08389% | 1192 | Contamination | 1
[20 empty] | 0 | 0% |  | Knob Collar Press Incorrectly
[9 empty] | 3 | [10 empty] | 3 | 0.2517% | 1192 | Incorrect Weld
[20 empty] | 0 | 0% | 1192 | Missing Harness | 0
[8 empty] | 4 | 1 | [10 empty] | 1 | 0.08389% | 1192 | Rough Knob Actuation
[20 empty] | 0 | 0% | 1192 | Squeaky Blade | 0 | BP Cap Separation:
[8 empty] | 4 | 2 | [9 empty] | 1 | 3 | 0.2517% | 1192 | Continuity Fail | 3 | Damaged Blade Pusher: 36ea
[20 empty] | 0 | 0% | 1192 | High Jaw Force | 0
[8 empty] | 1 | [11 empty] | 0 | 0% | 1192 | Low Jaw Force | 0
[20 empty] | 0 | 0% | 1192 | Return Spring Pop Out | 0
[8 empty] | 2 | [11 empty] | 0 | 0% | 1192 | Fuse Switch | 0
[20 empty] | 0 | 0% | 1192 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 1192 | Scratched Logo | 0
[8 empty] | 4 | [11 empty] | 0 | 0% | 1192 | Misassembled | 0
[20 empty] | 0 | 0% | 1192 | Trigger Actuation | 0
[8 empty] | 9 | [10 empty] | 2 | 2 | 0.1678% | 1192 | Cosmetic Defect | 2
[8 empty] | 1 | [11 empty] | 0 | 0% | 1192 | Fail DKP Program | 0
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 1192 | Cosmetic Defect | 0
[7 empty] | 2 | [12 empty] | 2 | 0.1678% | 1192 | Incorrect Jaw Crimp | 2 | 30ea Incorrect Hub Press
[7 empty] | 8 | [12 empty] | 8 | 0.6711% | 1192 | Insulation Damage | 8 | 32ea Damaged Pull Tube
[7 empty] | 1 | [12 empty] | 1 | 0.08389% | 1192 | Incorrect Weld | 1
[7 empty] | 2 | [12 empty] | 2 | 0.1678% | 1192 | Damaged Component | 2 | 7ea Damaged Upper Jaw
[7 empty] | 2 | [12 empty] | 2 | 0.1678% | 1192 | Stuck Blade | 2 | 2ea Damaged Lower Jaw
[7 empty] | 1 | [12 empty] | 1 | 0.08389% | 1192 | Exposed Wire | 1
[6 empty] | Total | 40 | 25 | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 21 | 102 | 8.557% | 1192 |  | 102

## EB012-EB212 Graphs
EB012/EB212 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 25
[5 empty] | 0.003 Jaw Gap Fail | 6
[5 empty] | Rough Actuation of Jaw | 4
[5 empty] | Damaged Harness | 2
[5 empty] | Jaw Short | 2
[5 empty] | Fuse Switch | 1
[5 empty] | Low Jaw Force | 1
[5 empty] | High Jaw Force | 1
[5 empty] | Front Stops Not Touching | 0
[5 empty] | Insulation Damage | 0
[5 empty] | Damaged Component | 0
[5 empty] | Continuity Fail | 0
[5 empty] | Incorrect Insulation Orientation | 0
[5 empty] | Misassembled | 0
[5 empty] | Damaged Jaw | 0
[5 empty] | Rough Knob Actuation | 0
Shop Order's Yield |  |  |  |  | Blade Activation | 0
Shop Order | Build QTY | Yield | Yield % | Date | Squeaky Blade | 0
1490055 | 590 | 537 | 91.02% | 2023-04-25 | Rough Trigger | 0
| #N/A | #N/A | #N/A | #N/A | Middle Stops Not Touching | 0
| #N/A | #N/A | #N/A | #N/A | Incorrect knob collar press | 0
| #N/A | #N/A | #N/A | #N/A | Both Not Touching | 0
| #N/A | #N/A | #N/A | #N/A | Loose Knob Collar | 0
| #N/A | #N/A | #N/A | #N/A | 0.005 Jaw Gap Fail | 0
Average Yield: |  |  | 91.02% |  | Damaged Handle | 0
[5 empty] | Discolored Jaws | 0

## EB012-EB212
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1490055 | EB212 | 576 | 590 | 537 | 91.02% | 2023-04-25 | [16 empty] | Top Level Non-Conformance | Total
[7 empty] | 1 |  | 4 | [9 empty] | 20 | 25 | 4.237% | 590 | Stuck Blade | 25
[7 empty] | 2 | [11 empty] | 2 | 4 | 0.678% | 590 | Rough Actuation of Jaw | 4
[7 empty] | 1 |  | 1 | [9 empty] | 4 | 6 | 1.017% | 590 | 0.003 Jaw Gap Fail | 6
[20 empty] | 0 | 0% | 590 | 0.006 Jaw Gap Fail | 0
[20 empty] | 0 | 0% | 590 | Front Stops Not Touching | 0
[20 empty] | 0 | 0% | 590 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 590 | Fuse Switch | 0
[20 empty] | 0 | 0% | 590 | Blade Does Not Pass Front Stop | 0
[9 empty] | 1 | [9 empty] | 1 | 2 | 0.339% | 590 | Jaw Short | 2
[20 empty] | 0 | 0% | 590 | Insulation Damage | 0
[20 empty] | 0 | 0% | 590 | Incorrect Insulation Orientation | 0
[20 empty] | 0 | 0% | 590 | Incorrect/Missing Weld | 0
[20 empty] | 0 | 0% | 590 | Knob Collar Press Incorrectly
[9 empty] | 2 | [10 empty] | 2 | 0.339% | 590 | Damaged Harness | 2
[20 empty] | 0 | 0% | 590 | Rough Knob Actuation
[20 empty] | 0 | 0% | 590 | Squeaky Blade | 0
[8 empty] | 2 | [11 empty] | 0 | 0% | 590 | Continuity Fail | 0
[19 empty] | 1 | 1 | 0.1695% | 590 | High Jaw Force | 1 | BP Cap Separation: ea
[8 empty] | 8 | 1 | [10 empty] | 1 | 0.1695% | 590 | Low Jaw Force | 1 | Damaged Blade Pusher : 130ea
[8 empty] | 1 | 1 | [10 empty] | 1 | 0.1695% | 590 | Return Spring Pop Out | 1 | Damaged Insulation : 50ea
[9 empty] | 1 | [10 empty] | 1 | 0.1695% | 590 | Fuse Switch | 1
[20 empty] | 0 | 0% | 590 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 590 | Scratched Logo | 0
[8 empty] | 5 | [11 empty] | 0 | 0% | 590 | Misassembled | 0
[20 empty] | 0 | 0% | 590 | Missing Component | 0
[20 empty] | 0 | 0% | 590 | Cosmetic Defect | 0
[20 empty] | 0 | 0% | 590 | Damaged Harness | 0
[23 empty] | Shafts | 0
[7 empty] | 1 | [12 empty] | 1 | 0.1695% | 590 | Cosmetic Defect | 1
[20 empty] | 0 | 0% | 590 | Incorrect Jaw Crimp | 0 | 2ea Incorrect Hub Press
[7 empty] | 5 | [12 empty] | 5 | 0.8475% | 590 | Contamination | 5 | 5ea Damaged Pull Tube
[20 empty] | 0 | 0% | 590 | Incorrect/Missing Weld | 0 | 9ea Damaged Upper Jaw
[7 empty] | 1 | [12 empty] | 1 | 0.1695% | 590 | Damaged Blade | 1 | 8ea Ddamgaed Cover Tube
[20 empty] | 0 | 0% | 590 | Blade Pusher Coming Out | 0 | 5ea Ddamgaed Lower Jaw
[7 empty] | 3 | [12 empty] | 3 | 0.5085% | 590 | Exposed Wire | 3
[6 empty] | Total | 14 | 16 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 28 | 53 | 8.983% | 590 |  | 53

## EB013-EB213 Graphs
EB012/EB212 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 108
[5 empty] | 0.003 Jaw Gap Fail | 40
[5 empty] | Incorrect Insulation Orientation | 21
[5 empty] | Damaged Harness | 16
[5 empty] | Insulation Damage | 14
[5 empty] | Rough Actuation of Jaw | 11
[5 empty] | Jaw Short | 11
[5 empty] | Fuse Switch | 8
[5 empty] | Continuity Fail | 8
[5 empty] | Front Stops Not Touching | 5
[5 empty] | Damaged Component | 3
[5 empty] | Misassembled | 2
[5 empty] | Low Jaw Force | 1
[5 empty] | Damaged Jaw | 0
[5 empty] | High Jaw Force | 0
[5 empty] | Rough Knob Actuation | 0
Shop Order's Yield |  |  |  |  | Blade Activation | 0
Shop Order | Build QTY | Yield | Yield % | Date | Squeaky Blade | 0
1486326 | 1255 | 1112 | 88.61% | 2023-04-06 | Rough Trigger | 0
1492704 | 1269 | 1118 | 88.1% | 2023-06-19 | Middle Stops Not Touching | 0
| #N/A | #N/A | #N/A | #N/A | Incorrect knob collar press | 0
| #N/A | #N/A | #N/A | #N/A | Both Not Touching | 0
| #N/A | #N/A | #N/A | #N/A | Loose Knob Collar | 0
| #N/A | #N/A | #N/A | #N/A | 0.005 Jaw Gap Fail | 0
Average Yield: |  |  | 88.35% |  | Damaged Handle | 0
[5 empty] | Discolored Jaws | 0

## EB013-EB213
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1486326 | EB213 | 1152 | 1255 | 1112 | 88.61% | 2023-04-06 | [16 empty] | Top Level Non-Conformance | Total
[7 empty] | 29 |  | 3 | [9 empty] | 10 | 42 | 3.347% | 1255 | Stuck Blade | 42
[7 empty] | 9 | [12 empty] | 9 | 0.7171% | 1255 | Rough Actuation of Jaw | 9
[7 empty] | 25 |  | 1 | [9 empty] | 5 | 31 | 2.47% | 1255 | 0.003 Jaw Gap Fail | 31
[20 empty] | 0 | 0% | 1255 | 0.006 Jaw Gap Fail | 0
[7 empty] | 2 |  | 1 | [10 empty] | 3 | 0.239% | 1255 | Front Stops Not Touching | 3
[20 empty] | 0 | 0% | 1255 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 1255 | Damaged Component | 0
[20 empty] | 0 | 0% | 1255 | Blade Does Not Pass Front Stop | 0
[7 empty] | 5 | [12 empty] | 5 | 0.3984% | 1255 | Jaw Short | 5
[20 empty] | 0 | 0% | 1255 | Insulation Damage | 0
[20 empty] | 0 | 0% | 1255 | Incorrect Insulation Orientation | 0
[20 empty] | 0 | 0% | 1255 | Missing Harness | 0
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 1255 | Incorrect Weld
[9 empty] | 4 | [10 empty] | 4 | 0.3187% | 1255 | Damaged Harness | 4
[8 empty] | 5 | 1 | [10 empty] | 1 | 0.07968% | 1255 | Rough Knob Actuation
[8 empty] | 1 | [11 empty] | 0 | 0% | 1255 | Squeaky Blade | 0
[8 empty] | 5 | [11 empty] | 0 | 0% | 1255 | Continuity Fail | 0
[8 empty] | 1 | [11 empty] | 0 | 0% | 1255 | High Jaw Force | 0 | BP Cap Separation: ea
[8 empty] | 4 | 1 | [10 empty] | 1 | 0.07968% | 1255 | Low Jaw Force | 1 | Damaged Blade Pusher : ea
[20 empty] | 0 | 0% | 1255 | Return Spring Pop Out | 0
[9 empty] | 6 | [10 empty] | 6 | 0.4781% | 1255 | Fuse Switch | 6
[8 empty] | 1 | [11 empty] | 0 | 0% | 1255 | Blade Lever Stop Pop Out | 0
[8 empty] | 1 | [10 empty] | 3 | 3 | 0.239% | 1255 | Scratched Logo | 3
[8 empty] | 13 | 1 | [10 empty] | 1 | 0.07968% | 1255 | Misassembled | 1
[8 empty] | 1 | 1 | [10 empty] | 1 | 0.07968% | 1255 | Missing Component | 1
[8 empty] | 1 | [9 empty] | 3 |  | 0 | 0% | 1255 | Cosmetic Defect | 0
[20 empty] | 0 | 0% | 1255 | Damaged Harness | 0
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 1255 | Incorrect Blade Stop Crimp | 0
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 1255 | Damaged Nut Thread | 1 | 3ea Incorrect Blade Stop Crimp (No Jaws)
[20 empty] | 0 | 0% | 1255 | Incorrect Blade Orientation | 0 | 2ea Damaged Upper Jaw
[7 empty] | 2 | [12 empty] | 2 | 0.1594% | 1255 | Incorrect Weld | 2 | 9ea Damaged Pull Tube
[7 empty] | 2 | [12 empty] | 2 | 0.1594% | 1255 | Insulation Damaged | 2
[7 empty] | 22 | [12 empty] | 22 | 1.753% | 1255 | Blade Pusher Coming Out | 22
[7 empty] | 5 | [12 empty] | 5 | 0.3984% | 1255 | Contamination | 5
[6 empty] | Total | 103 | 33 | 19 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 18 | 140 | 11.16% | 1255 |  | 140
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1489750 | EB213 | 40 | 40 | 37 | 92.5% | 2023-05-01 | [16 empty] | Top Level Non-Conformance | Total
[20 empty] | 0 | 0% | 40 | Stuck Blade | 0
[20 empty] | 0 | 0% | 40 | Rough Actuation of Jaw | 0
[20 empty] | 0 | 0% | 40 | 0.003 Jaw Gap Fail | 0
[20 empty] | 0 | 0% | 40 | 0.006 Jaw Gap Fail | 0
[20 empty] | 0 | 0% | 40 | Front Stops Not Touching | 0
[20 empty] | 0 | 0% | 40 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 40 | Damaged Component | 0
[20 empty] | 0 | 0% | 40 | Blade Does Not Pass Front Stop | 0
[20 empty] | 0 | 0% | 40 | Jaw Short | 0
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 40 | Insulation Damage | 1
[20 empty] | 0 | 0% | 40 | Incorrect Insulation Orientation | 0
[20 empty] | 0 | 0% | 40 | Missing Harness | 0
[20 empty] | 0 | 0% | 40 | Incorrect Weld
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 40 | Damaged Harness | 1
[20 empty] | 0 | 0% | 40 | Rough Knob Actuation
[20 empty] | 0 | 0% | 40 | Squeaky Blade | 0
[20 empty] | 0 | 0% | 40 | Continuity Fail | 0
[20 empty] | 0 | 0% | 40 | High Jaw Force | 0 | BP Cap Separation: ea
[20 empty] | 0 | 0% | 40 | Low Jaw Force | 0 | Damaged Blade Pusher : ea
[20 empty] | 0 | 0% | 40 | Return Spring Pop Out | 0
[20 empty] | 0 | 0% | 40 | Fuse Switch | 0
[20 empty] | 0 | 0% | 40 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 40 | Scratched Logo | 0
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 40 | Misassembled | 1
[20 empty] | 0 | 0% | 40 | Missing Component | 0
[20 empty] | 0 | 0% | 40 | Cosmetic Defect | 0
[20 empty] | 0 | 0% | 40 | Damaged Harness | 0
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 40 | Incorrect Blade Stop Crimp | 0
[20 empty] | 0 | 0% | 40 | Damaged Nut Thread | 0 | ea Incorrect Blade Stop Crimp (No Jaws)
[20 empty] | 0 | 0% | 40 | Incorrect Blade Orientation | 0 | Damaged Upper Jaw
[20 empty] | 0 | 0% | 40 | Incorrect Weld | 0 | ea Damaged Pull Tube
[20 empty] | 0 | 0% | 40 | Insulation Damaged | 0
[20 empty] | 0 | 0% | 40 | Blade Pusher Coming Out | 0
[20 empty] | 0 | 0% | 40 | Contamination | 0
[6 empty] | Total | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0.239% | 40 |  | 3
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1492704 | EB213 | 1152 | 1269 | 1118 | 88.1% | 2023-06-19 | [16 empty] | Top Level Non-Conformance | Total
[7 empty] | 56 |  | 3 | [9 empty] | 7 | 66 | 5.259% | 1269 | Stuck Blade | 66
[7 empty] | 2 | [12 empty] | 2 | 0.1594% | 1269 | Rough Actuation of Jaw | 2
[7 empty] | 8 | [11 empty] | 1 | 9 | 0.7171% | 1269 | 0.003 Jaw Gap Fail | 9
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 1269 | 0.006 Jaw Gap Fail | 1
[7 empty] | 2 | [12 empty] | 2 | 0.1594% | 1269 | Front Stops Not Touching | 2
[20 empty] | 0 | 0% | 1269 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 1269 | Damaged Component | 0
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 1269 | Blade Does Not Pass Front Stop | 1
[7 empty] | 4 | [11 empty] | 2 | 6 | 0.4781% | 1269 | Jaw Short | 6
[7 empty] | 11 |  | 1 | [9 empty] | 1 | 13 | 1.036% | 1269 | Insulation Damage | 13
[7 empty] | 20 |  | 1 | [10 empty] | 21 | 1.673% | 1269 | Incorrect Insulation Orientation | 21
[20 empty] | 0 | 0% | 1269 | Missing Harness | 0
[20 empty] | 0 | 0% | 1269 | Incorrect Weld
[7 empty] | 2 |  | 8 | [10 empty] | 10 | 0.7968% | 1269 | Damaged Harness | 10
[8 empty] | 1 | [11 empty] | 0 | 0% | 1269 | Rough Knob Actuation
[20 empty] | 0 | 0% | 1269 | Squeaky Blade | 0
[8 empty] | 2 | 1 | [9 empty] | 7 | 8 | 0.6375% | 1269 | Continuity Fail | 8
[8 empty] | 1 | [11 empty] | 0 | 0% | 1269 | High Jaw Force | 0 | BP Cap Separation: 10ea
[8 empty] | 1 | [11 empty] | 0 | 0% | 1269 | Low Jaw Force | 0 | Damaged Blade Pusher : 116ea
[20 empty] | 0 | 0% | 1269 | Return Spring Pop Out | 0
[8 empty] | 4 | 2 | [10 empty] | 2 | 0.1594% | 1269 | Fuse Switch | 2
[20 empty] | 0 | 0% | 1269 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 1269 | Scratched Logo | 0
[8 empty] | 24 | [11 empty] | 0 | 0% | 1269 | Misassembled | 0
[8 empty] | 1 | [11 empty] | 0 | 0% | 1269 | Missing Component | 0
[8 empty] | 2 | [11 empty] | 0 | 0% | 1269 | Cosmetic Defect | 0
[9 empty] | 1 | [10 empty] | 1 | 0.07968% | 1269 | Damaged Harness | 1
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 1269 | Incorrect Blade Stop Crimp | 0 | 1ea Contamination Upper Jaw
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 1269 | Jaw Won't Open | 1 | 5ea Incorrect Blade Stop Crimp (No Jaws)
[20 empty] | 0 | 0% | 1269 | Incorrect Blade Orientation | 0 | 2ea Incorrect Hub Press
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 1269 | Incorrect Weld | 1 | 7ea Damaged Pull Tube
[7 empty] | 1 | [12 empty] | 1 | 0.07968% | 1269 | Insulation Damaged | 1
[7 empty] | 3 | [12 empty] | 3 | 0.239% | 1269 | Damaged Component | 3
[7 empty] | 4 | [12 empty] | 4 | 0.3187% | 1269 | Exposed Wire | 4
[6 empty] | Total | 117 | 36 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 152 | 12.11% | 1269 |  | 152

## EB014-EB214 Graphs
EB012/EB212 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 17
[5 empty] | Damaged Harness | 5
[5 empty] | Fuse Switch | 2
[5 empty] | 0.003 Jaw Gap Fail | 1
[5 empty] | Low Jaw Force | 1
[5 empty] | Rough Actuation of Jaw | 0
[5 empty] | Jaw Short | 0
[5 empty] | Front Stops Not Touching | 0
[5 empty] | Misassembled | 0
[5 empty] | Insulation Damage | 0
[5 empty] | Damaged Component | 0
[5 empty] | Continuity Fail | 0
[5 empty] | Incorrect Insulation Orientation | 0
[5 empty] | Damaged Jaw | 0
[5 empty] | High Jaw Force | 0
[5 empty] | Rough Knob Actuation | 0
Shop Order's Yield |  |  |  |  | Blade Activation | 0
Shop Order | Build QTY | Yield | Yield % | Date | Squeaky Blade | 0
1490056 | 309 | 277 | 89.64% | 2023-04-25 | Rough Trigger | 0
| #N/A | #N/A | #N/A | #N/A | Middle Stops Not Touching | 0
| #N/A | #N/A | #N/A | #N/A | Incorrect knob collar press | 0
| #N/A | #N/A | #N/A | #N/A | Both Not Touching | 0
| #N/A | #N/A | #N/A | #N/A | Loose Knob Collar | 0
| #N/A | #N/A | #N/A | #N/A | 0.005 Jaw Gap Fail | 0
Average Yield: |  |  | 89.64% |  | Damaged Handle | 0
[5 empty] | Discolored Jaws | 0

## EB014-EB214
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | 5th Rework (New Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | Build Qty | Description | Total | Comments
1490056 | EB214 | 288 | 309 | 277 | 89.64% | 2023-04-25 | [16 empty] | Top Level Non-Conformance | Total
[7 empty] | 14 | [11 empty] | 3 | 17 | 5.502% | 309 | Stuck Blade | 17
[20 empty] | 0 | 0% | 309 | Rough Actuation of Jaw | 0
[7 empty] | 1 | [12 empty] | 1 | 0.3236% | 309 | 0.003 Jaw Gap Fail | 1
[20 empty] | 0 | 0% | 309 | 0.006 Jaw Gap Fail | 0
[20 empty] | 0 | 0% | 309 | Front Stops Not Touching | 0
[20 empty] | 0 | 0% | 309 | Middle Stops Not Touching | 0
[20 empty] | 0 | 0% | 309 | Damaged Component | 0
[20 empty] | 0 | 0% | 309 | Blade Does Not Pass Front Stop | 0
[20 empty] | 0 | 0% | 309 | Jaw Short | 0
[20 empty] | 0 | 0% | 309 | Insulation Damage | 0
[20 empty] | 0 | 0% | 309 | Incorrect Insulation Orientation | 0
[7 empty] | 1 | [11 empty] | 1 | 2 | 0.6472% | 309 | Fuse Switch | 2
[20 empty] | 0 | 0% | 309 | Incorrect Weld
[9 empty] | 5 | [10 empty] | 5 | 1.618% | 309 | Damaged Harness | 5
[9 empty] | 1 | [10 empty] | 1 | 0.3236% | 309 | Rough Knob Actuation
[20 empty] | 0 | 0% | 309 | Squeaky Blade | 0
[20 empty] | 0 | 0% | 309 | Continuity Fail | 0
[8 empty] | 1 | [11 empty] | 0 | 0% | 309 | High Jaw Force | 0 | BP Cap Separation: ea
[19 empty] | 1 | 1 | 0.3236% | 309 | Low Jaw Force | 1 | Damaged Blade Pusher : 16ea
[20 empty] | 0 | 0% | 309 | Return Spring Pop Out | 0
[20 empty] | 0 | 0% | 309 | Fuse Switch | 0
[20 empty] | 0 | 0% | 309 | Blade Lever Stop Pop Out | 0
[20 empty] | 0 | 0% | 309 | Scratched Logo | 0
[20 empty] | 0 | 0% | 309 | Misassembled | 0
[20 empty] | 0 | 0% | 309 | Missing Component | 0
[20 empty] | 0 | 0% | 309 | Cosmetic Defect | 0
[20 empty] | 0 | 0% | 309 | Visual | 0
[23 empty] | Shafts | 0
[20 empty] | 0 | 0% | 309 | Incorrect Blade Stop Crimp | 0
[20 empty] | 0 | 0% | 309 | Incorrect Jaw Crimp | 0 | ea Incorrect Blade Stop Crimp (No Jaws)
[20 empty] | 0 | 0% |  | Incorrect Blade Orientation |  | 14ea Contamination Blade
[20 empty] | 0 | 0% |  | Incorrect Weld
[20 empty] | 0 | 0% |  | Damaged Component
[7 empty] | 1 | [12 empty] | 1 | 0.3236% | 309 | Contamination | 1
[20 empty] | 0 | 0% | 309 | Defective Component | 0
[20 empty] | 0 | 0% | 309 | Damaged Component | 0
[7 empty] | 1 | [12 empty] | 1 | 0.3236% | 309 | Blacde Pusher Coming Out | 1
[7 empty] | 3 | [12 empty] | 3 | 0.9709% | 309 | Exposed Wire | 3
[6 empty] | Total | 21 | 1 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 32 | 10.36% | 309 |  | 32

## EB030-EB230 Graphs
EB030/EB230 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 216
[5 empty] | Front Stop not Touching | 22
[5 empty] | Continuity Fail | 20
[5 empty] | Discolored Weld | 20
[5 empty] | 0.003 Jaw Gap Fail | 13
[5 empty] | 0.007 Jaw Gap Fail | 12
[5 empty] | Exposed Wire | 9
[5 empty] | Jaw Isolation | 6
[5 empty] | Damaged Harness | 4
[5 empty] | Fuse Switch | 2
[5 empty] | Cosmetic Defect | 1
[5 empty] | Electrode Isolation | 0
[5 empty] | Low Jaw Force | 0
[5 empty] | High Jaw Force | 0
[5 empty] | Contamination | 0
[5 empty] | Blade Activation | 0
Shop Order's Yield |  |  |  |  | Engineering Scrap | 0
Shop Order | Build QTY | Yield | Yield % | Date
1484680 | 1956 | 1775 | 90.75% | 2023-04-23
1486766 | 2534 | 2326 | 91.79% | 2023-05-05
1490572 | 601 | 539 | 89.68% | 2023-05-09
1492604 | 506 | 461 | 91.11% | 2023-06-01
1488012 | 1963 | 1782 | 90.78% | 2023-06-14
1489769 | 1972 | 1770 | 89.76% | 2023-06-26
Average Yield: |  |  | 90.64%
Formulas:
Date = VLOOKUP(Table14310[[#This Row],[Shop Order]],'EB030-EB230'!A:AE,8,FALSE)

## EB030-EB230
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1484679 | EB230 | 1920 | 1948 | 1814 | 93.12% | 4.363% | 2023-04-01 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[8 empty] | 1 |  |  | 1 | 0.05133% | 1948 | 0.003 Jaw Gap Fail | 1
[8 empty] | 7 |  |  | 7 | 0.3593% | 1948 | 0.007 Jaw Gap Fail | 7
[11 empty] | 0 | 0% | 1948 | High Jaw Force | 0
[11 empty] | 0 | 0% | 1948 | Low Jaw Force | 0
[8 empty] | 73 |  | 11 | 84 | 4.312% | 1948 | Stuck Blade | 84
[11 empty] | 0 | 0% | 1948 | Blade Activation | 0
[11 empty] | 0 | 0% | 1948 | Electrode Isolation | 0
[8 empty] | 1 |  |  | 1 | 0.05133% | 1948 | Jaw Isolation | 1
[10 empty] | 1 | 1 | 0.05133% | 1948 | Continuity Fail | 1
[11 empty] | 0 | 0% | 1948 | Damaged Harness | 0
[8 empty] | 7 |  |  | 7 | 0.3593% | 1948 | Missing Harness | 7
[11 empty] | 0 | 0% | 1948 | Missing Front Stop Dot | 0
[11 empty] | 0 | 0% | 1948 | Desired Pattern Could Not Be Found | 0
[11 empty] | 0 | 0% | 1948 | Blade Does Not Pass Front Stop | 0
[11 empty] | 0 | 0% | 1948 | Front Stop not Touching | 0
[8 empty] | 2 |  |  | 2 | 0.1027% | 1948 | Cosmetic Defect | 2
[8 empty] | 2 |  |  | 2 | 0.1027% | 1948 | Exposed Wire | 2
[11 empty] | 0 | 0% | Build QTY | Incomplete / Incorrect Weld | 0
[8 empty] | 3 |  |  | 3 | 0.154% | 1948 | Damaged Component | 3 | Jaw
[14 empty] | Reworkables | 0
[9 empty] | 13 |  | 0 | 0% | 1948 | Blade Link Pop Out | 0
[9 empty] | 9 | 5 | 5 | 0.2567% | 1948 | Scratched Logo | 5
[11 empty] | 0 | 0% | 1948 | Cosmetic Defect | 0
[9 empty] | 13 |  | 0 | 0% | 1948 | Missing Component | 0 | 8ea No Grease
[9 empty] | 1 |  | 0 | 0% | 1948 | Particulate Matter | 0
[9 empty] | 28 |  | 0 | 0% | 1948 | Misassembled Component | 0
[9 empty] | 7 |  | 0 | 0% | 1948 | Wire Popped Out | 0
[9 empty] | 14 |  | 0 | 0% | 1948 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[11 empty] | 0 | 0% | Build QTY | Damaged Component (Static) | 0
|  |  | . |  |  |  |  | 1 |  |  | 0 | 0% | Build QTY | Incomplete / Incorrect Weld (Static) | 0
[11 empty] | 0 | 0% | 1948 | Cosmetic Defect | 0 | Scratched Logo :
[8 empty] | 14 |  |  | 0 | 0% | Build QTY | Exposed Wire ( Dynamic) | 0 | 104ea ARM COVER, STATIC
[8 empty] | 2 |  |  | 0 | 0% | 1948 | Exposed Wire (Static) | 0 | 291ea ARM HOUSING, STATIC
[8 empty] | 4 |  |  | 4 | 0.2053% | 1948 | Exposed Wire | 4
[8 empty] | 1 |  |  | 1 | 0.05133% | 1948 | Cosmetic Defect | 1
[8 empty] | 12 |  |  | 12 | 0.616% | 1948 | Front Stop not Touching | 12
[8 empty] | 2 |  |  | 2 | 0.1027% | 1948 | Blade Guide Does Not Go Through | 2
[11 empty] | 0 | 0% | Build QTY | Discolored Weld | 0
[8 empty] | 6 |  |  | 6 | 0.308% | 1948 | Incomplete / Incorrect Weld | 6
[8 empty] | 3 |  |  | 3 | 0.154% | 1948 | Damaged Component | 3 | Bend Jaw
[8 empty] | 2 |  |  | 0 | 0% | 1948 | Contamination ( Dynamic) | 0
[11 empty] | 0 | 0% | 1948 | Contamination | 0
[8 empty] | 1 |  |  | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[8 empty] | 4 |  |  | 0 | 0% | 1948 | OverSize Shaft | 0
[7 empty] | Total | 148 | 85 | 17 | 141 | 7.238% | 1948
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1486788 | EB230 | 360 | 367 | 330 | 89.92% | 6.54% | 2023-04-06 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[11 empty] | 0 | 0% | 367 | 0.003 Jaw Gap Fail | 0
[8 empty] | 2 |  |  | 2 | 0.545% | 367 | 0.007 Jaw Gap Fail | 2
[11 empty] | 0 | 0% | 367 | High Jaw Force | 0
[11 empty] | 0 | 0% | 367 | Low Jaw Force | 0
[8 empty] | 19 |  |  | 19 | 5.177% | 367 | Stuck Blade | 19
[11 empty] | 0 | 0% | 367 | Blade Activation | 0
[8 empty] | 6 |  |  | 6 | 1.635% | 367 | Focus Error | 6
[11 empty] | 0 | 0% | 367 | Jaw Isolation | 0
[11 empty] | 0 | 0% | 367 | Continuity Fail | 0
[11 empty] | 0 | 0% | 367 | Damaged Harness | 0
[11 empty] | 0 | 0% | 367 | Missing Harness | 0
[11 empty] | 0 | 0% | 367 | Missing Front Stop Dot | 0
[11 empty] | 0 | 0% | 367 | Desired Pattern Could Not Be Found | 0
[11 empty] | 0 | 0% | 367 | Blade Does Not Pass Front Stop | 0
[11 empty] | 0 | 0% | 367 | Front Stop not Touching | 0
[8 empty] | 1 |  |  | 1 | 0.2725% | 367 | Cosmetic Defect | 1
[11 empty] | 0 | 0% | 367 | Exposed Wire | 0
[11 empty] | 0 | 0% | Build QTY | Incomplete / Incorrect Weld | 0
[8 empty] | 2 |  |  | 2 | 0.545% | 367 | Damaged Component | 2 | Bend Jaw
[14 empty] | Reworkables | 0
[9 empty] | 3 |  | 0 | 0% | 367 | Blade Link Pop Out | 0
[9 empty] | 1 |  | 0 | 0% | 367 | Scratched Logo | 0
[11 empty] | 0 | 0% | 367 | Cosmetic Defect | 0
[11 empty] | 0 | 0% | 367 | Missing Component | 0 | ea No Grease
[9 empty] | 5 |  | 0 | 0% | 367 | Particulate Matter | 0
[9 empty] | 14 |  | 0 | 0% | 367 | Misassembled Component | 0
[9 empty] | 1 |  | 0 | 0% | 367 | Wire Popped Out | 0
[11 empty] | 0 | 0% | 367 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[11 empty] | 0 | 0% | Build QTY | Damaged Component (Static) | 0
|  |  | . | [7 empty] | 0 | 0% | Build QTY | Incomplete / Incorrect Weld (Static) | 0
[11 empty] | 0 | 0% | 367 | Cosmetic Defect | 0 | Scratched Logo :
[11 empty] | 0 | 0% | Build QTY | Exposed Wire ( Dynamic) | 0 | ea ARM COVER, STATIC
[11 empty] | 0 | 0% | 367 | Exposed Wire (Static) | 0 | ea ARM HOUSING, STATIC
[8 empty] | 5 |  |  | 5 | 1.362% | 367 | Exposed Wire | 5
[11 empty] | 0 | 0% | 367 | Cosmetic Defect | 0
[11 empty] | 0 | 0% | 367 | Front Stop not Touching | 0
[11 empty] | 0 | 0% | 367 | Blade Guide Does Not Go Through | 0
[11 empty] | 0 | 0% | Build QTY | Discolored Weld | 0
[11 empty] | 0 | 0% | 367 | Incomplete / Incorrect Weld | 0
[8 empty] | 2 |  |  | 2 | 0.545% | 367 | Damaged Jaw | 2 | Bend Jaw
[11 empty] | 0 | 0% | 367 | Contamination ( Dynamic) | 0
[11 empty] | 0 | 0% | 367 | Contamination | 0
[11 empty] | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[11 empty] | 0 | 0% | 367 | OverSize Shaft | 0
[7 empty] | Total | 37 | 24 | 0 | 37 | 10.08% | 367
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1486224 | EB230 | 1920 | 1971 | 1722 | 87.37% | 4.921% | 2023-04-15 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[11 empty] | 0 | 0% | 1971 | 0.003 Jaw Gap Fail | 0
[8 empty] | 13 |  |  | 13 | 0.6596% | 1971 | 0.007 Jaw Gap Fail | 13
[11 empty] | 0 | 0% | 1971 | High Jaw Force | 0
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Low Jaw Force | 2
[8 empty] | 98 |  | 13 | 111 | 5.632% | 1971 | Stuck Blade | 111
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Wire Cut Short | 2
[8 empty] | 5 |  |  | 5 | 0.2537% | 1971 | Focus Error | 5
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Jaw Isolation | 2
[8 empty] | 1 |  | 1 | 2 | 0.1015% | 1971 | Continuity Fail | 2
[11 empty] | 0 | 0% | 1971 | Damaged Harness | 0
[8 empty] | 1 |  |  | 1 | 0.05074% | 1971 | Excess Grivory | 1
[8 empty] | 30 |  |  | 30 | 1.522% | 1971 | Missing Front/Back Stop Dot | 30
[8 empty] | 6 |  |  | 6 | 0.3044% | 1971 | Desired Pattern Could Not Be Found | 6
[11 empty] | 0 | 0% | 1971 | Blade Does Not Pass Front Stop | 0
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Front Stop not Touching | 2
[8 empty] | 4 |  |  | 4 | 0.2029% | 1971 | Cosmetic Defect | 4
[8 empty] | 1 |  |  | 1 | 0.05074% | 1971 | Exposed Wire | 1
[11 empty] | 0 | 0% | Build QTY | Incomplete / Incorrect Weld | 0
[8 empty] | 13 |  |  | 13 | 0.6596% | 1971 | Damaged Component | 13 | Bend Jaw
[14 empty] | Reworkables | 0
[9 empty] | 10 |  | 0 | 0% | 1971 | Blade Link Pop Out | 0
[9 empty] | 7 | 4 | 4 | 0.2029% | 1971 | Scratched Logo | 4
[9 empty] | 4 |  | 0 | 0% | 1971 | Cosmetic Defect | 0
[9 empty] | 8 |  | 0 | 0% | 1971 | Missing Component | 0 | 4ea No Grease
[9 empty] | 4 |  | 0 | 0% | 1971 | Particulate Matter | 0
[9 empty] | 50 |  | 0 | 0% | 1971 | Misassembled Component | 0
[9 empty] | 1 |  | 0 | 0% | 1971 | Wire Popped Out | 0
[9 empty] | 13 |  | 0 | 0% | 1971 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[8 empty] | 1 |  |  | 0 | 0% | Build QTY | Cosmetic Defect (Static) | 0
|  |  | . | [7 empty] | 0 | 0% | Build QTY | Incomplete / Incorrect Weld (Static) | 0
[8 empty] | 8 |  |  | 8 | 0.4059% | 1971 | Cosmetic Defect (Jaw) | 8 | Scratched Logo :
[8 empty] | 2 |  |  | 0 | 0% | Build QTY | Exposed Wire ( Dynamic) | 0 | 87ea ARM COVER, STATIC
[8 empty] | 1 |  |  | 0 | 0% | 1971 | Exposed Wire (Static) | 0 | 145ea ARM HOUSING, STATIC
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Exposed Wire | 2
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Defective Component (Jaw) | 2
[8 empty] | 28 |  |  | 28 | 1.421% | 1971 | Front Stop not Touching | 28
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Blade Guide Does Not Go Through | 2
[11 empty] | 0 | 0% | Build QTY | Discolored Weld | 0
[8 empty] | 4 |  |  | 4 | 0.2029% | 1971 | Incomplete / Incorrect Weld | 4
[8 empty] | 2 |  |  | 2 | 0.1015% | 1971 | Damaged Jaw | 2 | Bend Jaw
[11 empty] | 0 | 0% | 1971 | Contamination ( Dynamic) | 0
[8 empty] | 3 |  |  | 3 | 0.1522% | 1971 | Missing Front/Back Stop Dot | 3
[11 empty] | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[11 empty] | 0 | 0% | 1971 | OverSize Shaft | 0
[7 empty] | Total | 235 | 97 | 18 | 249 | 12.63% | 1971
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1484680 | EB230 | 1920 | 1956 | 1775 | 90.75% | 2.249% | 2023-04-23 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[8 empty] | 2 |  |  | 2 | 0.1022% | 1956 | 0.003 Jaw Gap Fail | 2
[8 empty] | 29 |  |  | 29 | 1.483% | 1956 | 0.007 Jaw Gap Fail | 29
[11 empty] | 0 | 0% | 1956 | High Jaw Force | 0
[10 empty] | 1 | 1 | 0.05112% | 1956 | Low Jaw Force | 1
[8 empty] | 56 |  | 10 | 66 | 3.374% | 1956 | Stuck Blade | 66
[8 empty] | 1 |  |  | 1 | 0.05112% | 1956 | Electrode Isolation | 1
[8 empty] | 9 |  |  | 9 | 0.4601% | 1956 | Focus Error | 9
[8 empty] | 2 |  |  | 2 | 0.1022% | 1956 | Jaw Isolation | 2
[8 empty] | 7 |  |  | 7 | 0.3579% | 1956 | Continuity Fail | 7
[8 empty] | 7 |  |  | 7 | 0.3579% | 1956 | Damaged Harness | 7
[8 empty] | 1 |  |  | 1 | 0.05112% | 1956 | Engineering Sample | 1
[8 empty] | 2 |  |  | 2 | 0.1022% | 1956 | Missing Front/Back Stop Dot | 2
[8 empty] | 7 |  |  | 7 | 0.3579% | 1956 | Desired Pattern Could Not Be Found | 7
[11 empty] | 0 | 0% | 1956 | Blade Does Not Pass Front Stop | 0
[8 empty] | 2 |  |  | 2 | 0.1022% | 1956 | Front Stop not Touching | 2
[11 empty] | 0 | 0% | 1956 | Cosmetic Defect | 0
[11 empty] | 0 | 0% | 1956 | Exposed Wire | 0
[11 empty] | 0 | 0% | Build QTY | Incomplete / Incorrect Weld | 0
[8 empty] | 4 |  |  | 4 | 0.2045% | 1956 | Damaged Component | 4 | Bend Jaw
[14 empty] | Reworkables | 0
[9 empty] | 7 |  | 0 | 0% | 1956 | Blade Link Pop Out | 0
[9 empty] | 4 |  | 0 | 0% | 1956 | Scratched Logo | 0
[9 empty] | 5 |  | 0 | 0% | 1956 | Cosmetic Defect | 0
[9 empty] | 8 |  | 0 | 0% | 1956 | Missing Component | 0 | 3ea No Grease
[9 empty] | 3 |  | 0 | 0% | 1956 | Particulate Matter | 0
[9 empty] | 13 |  | 0 | 0% | 1956 | Misassembled Component | 0
[11 empty] | 0 | 0% | 1956 | Wire Popped Out | 0
[9 empty] | 4 |  | 0 | 0% | 1956 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[8 empty] | 1 |  |  | 0 | 0% | Build QTY | Ben Jaw (Dynamic) | 0
|  |  | . | [7 empty] | 0 | 0% | Build QTY | Incomplete / Incorrect Weld (Static) | 0
[11 empty] | 0 | 0% | 1956 | Cosmetic Defect (Jaw) | 0 | Scratched Logo :
[8 empty] | 1 |  |  | 0 | 0% | Build QTY | Exposed Wire ( Dynamic) | 0 | ea ARM COVER, STATIC
[11 empty] | 0 | 0% | 1956 | Exposed Wire (Static) | 0 | ea ARM HOUSING, STATIC
[11 empty] | 0 | 0% | 1956 | Exposed Wire | 0
[11 empty] | 0 | 0% | 1956 | Defective Component (Jaw) | 0
[8 empty] | 25 |  |  | 25 | 1.278% | 1956 | Front Stop not Touching | 25
[8 empty] | 1 |  |  | 1 | 0.05112% | 1956 | Blade Guide Does Not Go Through | 1
[8 empty] | 1 |  |  | 1 | 0.05112% | Build QTY | Misassembled Component | 1
[8 empty] | 3 |  |  | 3 | 0.1534% | 1956 | Incomplete / Incorrect Weld | 3
[8 empty] | 1 |  |  | 1 | 0.05112% | 1956 | Damaged Jaw | 1 | Bend Jaw
[8 empty] | 3 |  |  | 0 | 0% | 1956 | Contamination ( Static) | 0
[8 empty] | 5 |  |  | 5 | 0.2556% | 1956 | Missing Front/Back Stop Dot | 5
[11 empty] | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[8 empty] | 2 |  |  | 0 | 0% | 1956 | OverSize Shaft | 0
[7 empty] | Total | 172 | 44 | 11 | 176 | 8.998% | 1956
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1486766 | EB230 | 2478 | 2534 | 2326 | 91.79% | 2.21% | 2023-05-05 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[8 empty] | 1 |  |  | 1 | 0.03946% | 2534 | 0.003 Jaw Gap Fail | 1
[8 empty] | 13 |  |  | 13 | 0.513% | 2534 | 0.007 Jaw Gap Fail | 13
[11 empty] | 0 | 0% | 2534 | High Jaw Force | 0
[11 empty] | 0 | 0% | 2534 | Low Jaw Force | 0
[8 empty] | 75 |  | 10 | 85 | 3.354% | 2534 | Stuck Blade | 85
[10 empty] | 2 | 2 | 0.07893% | 2534 | Electrode Isolation | 2
[8 empty] | 5 |  |  | 5 | 0.1973% | 2534 | Focus Error | 5
[11 empty] | 0 | 0% | 2534 | Jaw Isolation | 0
[8 empty] | 5 |  |  | 5 | 0.1973% | 2534 | Continuity Fail | 5
[8 empty] | 4 |  |  | 4 | 0.1579% | 2534 | Damaged Harness | 4
[8 empty] | 2 |  |  | 2 | 0.07893% | 2534 | Excess Grivory | 2
[8 empty] | 4 |  |  | 4 | 0.1579% | 2534 | Missing Front/Back Stop Dot | 4
[8 empty] | 1 |  |  | 1 | 0.03946% | 2534 | Particulate Matter | 1
[8 empty] | 6 |  |  | 6 | 0.2368% | 2534 | Fuse Switch | 6 | Taken By Engineering
[8 empty] | 1 |  |  | 1 | 0.03946% | 2534 | Exposed Wire | 1
[8 empty] | 5 |  | 7 | 12 | 0.4736% | 2534 | Cosmetic Defect | 12
[8 empty] | 3 |  |  | 3 | 0.1184% | 2534 | Pinch Cable | 3
[8 empty] | 6 |  |  | 6 | 0.2368% | Build QTY | Device Key Programming Fail | 6
[8 empty] | 2 |  |  | 2 | 0.07893% | 2534 | Damaged Component | 2 | Bend Jaw
[14 empty] | Reworkables | 0
[9 empty] | 8 |  | 0 | 0% | 2534 | Blade Link Pop Out | 0
[9 empty] | 9 |  | 0 | 0% | 2534 | Scratched Logo | 0
[9 empty] | 1 |  | 0 | 0% | 2534 | Cosmetic Defect | 0
[9 empty] | 5 |  | 0 | 0% | 2534 | Missing Component | 0 | 4ea No Grease
[9 empty] | 5 |  | 0 | 0% | 2534 | Particulate Matter | 0
[9 empty] | 24 |  | 0 | 0% | 2534 | Misassembled Component | 0 | 2ea Open Handle
[9 empty] | 4 |  | 0 | 0% | 2534 | Wire Popped Out | 0
[11 empty] | 0 | 0% | 2534 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[8 empty] | 2 |  |  | 0 | 0% | Build QTY | Ben Jaw (Dynamic) | 0
|  |  | . |  |  |  |  | 4 |  |  | 0 | 0% | Build QTY | Damaged omponent (Static) | 0
[8 empty] | 3 |  |  | 3 | 0.1184% | 2534 | Cosmetic Defect | 3 | Scratched Logo :
[8 empty] | 2 |  |  | 2 | 0.07893% | Build QTY | Wire Cut Short | 2 | 57ea ARM COVER, STATIC
[8 empty] | 1 |  |  | 0 | 0% | 2534 | Exposed Wire (Static) | 0 | 203ea ARM HOUSING, STATIC
[8 empty] | 12 |  |  | 12 | 0.4736% | 2534 | Exposed Wire | 12
[8 empty] | 3 |  |  | 3 | 0.1184% | 2534 | Damaged Component (Bend Jaw) | 3
[8 empty] | 23 |  |  | 23 | 0.9077% | 2534 | Front Stop not Touching | 23
[8 empty] | 4 |  |  | 4 | 0.1579% | 2534 | Blade Guide Does Not Go Through | 4
[8 empty] | 2 |  |  | 2 | 0.07893% | Build QTY | Misassembled Component | 2 | 52ea Discoloration Shaft
[8 empty] | 2 |  |  | 2 | 0.07893% | 2534 | Incomplete / Incorrect Weld | 2 | 28ea Done Welded Scrap
[8 empty] | 3 |  |  | 3 | 0.1184% | 2534 | Contamination | 3
[8 empty] | 1 |  |  | 0 | 0% | 2534 | Contamination ( Dynamic) | 0
[8 empty] | 2 |  |  | 2 | 0.07893% | 2534 | Missing Front/Back Stop Dot | 2
[8 empty] | 2 |  |  | 0 | 0% | 588 | Excess Grivory (Static) | 0
[11 empty] | 0 | 0% | 2534 | Discoloreation Shaft | 0
[7 empty] | Total | 199 | 56 | 19 | 208 | 8.208% | 2534
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1490572 | EB230 | 576 | 601 | 539 | 89.68% | 2.496% | 2023-05-09 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[8 empty] | 2 |  |  | 2 | 0.3328% | 601 | 0.003 Jaw Gap Fail | 2
[8 empty] | 2 |  |  | 2 | 0.3328% | 601 | 0.007 Jaw Gap Fail | 2
[11 empty] | 0 | 0% | 601 | High Jaw Force | 0
[11 empty] | 0 | 0% | 601 | Low Jaw Force | 0
[8 empty] | 28 |  |  | 28 | 4.659% | 601 | Stuck Blade | 28
[11 empty] | 0 | 0% | 601 | Electrode Isolation | 0
[11 empty] | 0 | 0% | 601 | Focus Error | 0
[11 empty] | 0 | 0% | 601 | Jaw Isolation | 0
[11 empty] | 0 | 0% | 601 | Continuity Fail | 0
[8 empty] | 1 |  |  | 1 | 0.1664% | 601 | Damaged Harness | 1
[11 empty] | 0 | 0% | 601 | Excess Grivory | 0
[8 empty] | 1 |  |  | 1 | 0.1664% | 601 | Missing Front/Back Stop Dot | 1
[11 empty] | 0 | 0% | 601 | Desired Pattern Could Not Be Found | 0
[11 empty] | 0 | 0% | 601 | Fuse Switch | 0 | Taken By Engineering
[11 empty] | 0 | 0% | 601 | Exposed Wire | 0
[8 empty] | 1 |  |  | 1 | 0.1664% | 601 | Cosmetic Defect | 1
[11 empty] | 0 | 0% | 601 | Pinch Cable | 0
[11 empty] | 0 | 0% | Build QTY | Divice Key Programming Fail | 0
[11 empty] | 0 | 0% | 601 | Damaged Component | 0 | Bend Jaw
[14 empty] | Reworkables | 0
[9 empty] | 3 |  | 0 | 0% | 601 | Blade Link Pop Out | 0
[9 empty] | 2 | 2 | 2 | 0.3328% | 601 | Scratched Logo | 2
[11 empty] | 0 | 0% | 601 | Cosmetic Defect | 0
[9 empty] | 3 |  | 0 | 0% | 601 | Missing Component | 0 | 2ea No Grease
[11 empty] | 0 | 0% | 601 | Particulate Matter | 0
[9 empty] | 5 |  | 0 | 0% | 601 | Misassembled Component | 0 | 0ea Open Handle
[9 empty] | 2 |  | 0 | 0% | 601 | Wire Popped Out | 0
[11 empty] | 0 | 0% | 601 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[11 empty] | 0 | 0% | Build QTY | Ben Jaw (Dynamic) | 0
|  |  | . | [7 empty] | 0 | 0% | Build QTY | Damaged omponent (Static) | 0
[8 empty] | 5 |  |  | 5 | 0.8319% | 601 | Cosmetic Defect | 5 | Scratched Logo :
[11 empty] | 0 | 0% | Build QTY | Wire Cut Short | 0 | 60ea ARM COVER, STATIC
[8 empty] | 2 |  |  | 0 | 0% | 601 | Exposed Wire (Static) | 0 | 31ea ARM HOUSING, STATIC
[8 empty] | 5 |  |  | 5 | 0.8319% | 601 | Exposed Wire | 5
[8 empty] | 2 |  |  | 2 | 0.3328% | 601 | Damaged Component (Bend Jaw) | 2
[8 empty] | 11 |  |  | 11 | 1.83% | 601 | Front Stop not Touching | 11
[8 empty] | 1 |  |  | 1 | 0.1664% | 601 | Blade Guide Does Not Go Through | 1
[11 empty] | 0 | 0% | Build QTY | Misassembled Component | 0
[8 empty] | 1 |  |  | 1 | 0.1664% | 601 | Incomplete / Incorrect Weld | 1
[11 empty] | 0 | 0% | 601 | Contamination | 0
[8 empty] | 1 |  |  | 0 | 0% | 601 | Incomplete / Incorrect Weld (Static) | 0
[8 empty] | 1 |  |  | 0 | 0% | 601 | Missing Front/Back Stop Dot (Static) | 0
[8 empty] | 1 |  |  | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[8 empty] | 3 |  |  | 0 | 0% | 601 | Undersized Shaft | 0
[7 empty] | Total | 68 | 15 | 2 | 62 | 2.447%
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1492604 | EB230 | 480 | 506 | 461 | 91.11% | 3.755% | 2023-06-01 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[8 empty] | 1 |  |  | 1 | 0.1976% | 506 | 0.003 Jaw Gap Fail | 1
[8 empty] | 2 |  |  | 2 | 0.3953% | 506 | 0.007 Jaw Gap Fail | 2
[11 empty] | 0 | 0% | 506 | High Jaw Force | 0
[11 empty] | 0 | 0% | 506 | Low Jaw Force | 0
[8 empty] | 8 |  | 3 | 11 | 2.174% | 506 | Stuck Blade | 11
[11 empty] | 0 | 0% | 506 | Electrode Isolation | 0
[11 empty] | 0 | 0% | 506 | Focus Error | 0
[11 empty] | 0 | 0% | 506 | Jaw Isolation | 0
[8 empty] | 1 |  |  | 1 | 0.1976% | 506 | Continuity Fail | 1
[8 empty] | 2 |  |  | 2 | 0.3953% | 506 | Damaged Harness | 2
[11 empty] | 0 | 0% | 506 | Excess Grivory | 0
[11 empty] | 0 | 0% | 506 | Missing Front/Back Stop Dot | 0
[11 empty] | 0 | 0% | 506 | Desired Pattern Could Not Be Found | 0
[11 empty] | 0 | 0% | 506 | Fuse Switch | 0
[11 empty] | 0 | 0% | 506 | Exposed Wire | 0
[11 empty] | 0 | 0% | 506 | Cosmetic Defect | 0
[11 empty] | 0 | 0% | 506 | Pinch Cable | 0
[8 empty] | 1 |  |  | 1 | 0.1976% | Build QTY | Contamination | 1
[11 empty] | 0 | 0% | 506 | Damaged Component | 0
[14 empty] | Reworkables | 0
[9 empty] | 7 |  | 0 | 0% | 506 | Blade Link Pop Out | 0
[11 empty] | 0 | 0% | 506 | Scratched Logo | 0
[11 empty] | 0 | 0% | 506 | Cosmetic Defect | 0
[9 empty] | 2 |  | 0 | 0% | 506 | Missing Component | 0 | 2ea No Grease
[9 empty] | 3 |  | 0 | 0% | 506 | Particulate Matter | 0
[9 empty] | 6 |  | 0 | 0% | 506 | Misassembled Component | 0
[9 empty] | 1 |  | 0 | 0% | 506 | Wire Popped Out | 0
[11 empty] | 0 | 0% | 506 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[8 empty] | 29 |  |  | 0 | 0% | Build QTY | Excess Grivory (Static) | 0
|  |  | . | [7 empty] | 0 | 0% | Build QTY | Damaged omponent (Static) | 0
[8 empty] | 1 |  |  | 1 | 0.1976% | 506 | Cosmetic Defect | 1 | Scratched Logo :
[11 empty] | 0 | 0% | Build QTY | Wire Cut Short | 0 | ea ARM COVER, STATIC
[11 empty] | 0 | 0% | 506 | Exposed Wire (Static) | 0 | ea ARM HOUSING, STATIC
[8 empty] | 1 |  |  | 1 | 0.1976% | 506 | Exposed Wire | 1
[11 empty] | 0 | 0% | 506 | Damaged Component (Bend Jaw) | 0
[11 empty] | 0 | 0% | 506 | Front Stop not Touching | 0
[8 empty] | 23 |  |  | 23 | 4.545% | 506 | Blade Guide Does Not Go Through | 23
[11 empty] | 0 | 0% | Build QTY | Misassembled Component | 0
[10 empty] | 1 | 1 | 0.1976% | 506 | Incomplete / Incorrect Weld | 1
[11 empty] | 0 | 0% | 506 | Contamination | 0
[11 empty] | 0 | 0% | 506 | Incomplete / Incorrect Weld (Static) | 0
[11 empty] | 0 | 0% | 506 | Missing Front/Back Stop Dot (Static) | 0
[11 empty] | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[11 empty] | 0 | 0% | 506 | Undersized Shaft | 0
[7 empty] | Total | 69 | 19 | 4 | 44 | 8.696%
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1488012 | EB230 | 1920 | 1963 | 1782 | 90.78% | 4.33% | 2023-06-14 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[8 empty] | 5 |  |  | 5 | 0.2547% | 1963 | 0.003 Jaw Gap Fail | 5
[8 empty] | 5 |  | 1 | 6 | 0.3057% | 1963 | 0.007 Jaw Gap Fail | 6
[11 empty] | 0 | 0% | 1963 | High Jaw Force | 0
[11 empty] | 0 | 0% | 1963 | Low Jaw Force | 0
[8 empty] | 85 |  | 14 | 99 | 5.043% | 1963 | Stuck Blade | 99
[11 empty] | 0 | 0% | 1963 | Electrode Isolation | 0
[11 empty] | 0 | 0% | 1963 | Focus Error | 0
[8 empty] | 2 |  | 2 | 4 | 0.2038% | 1963 | Jaw Isolation | 4
[8 empty] | 1 |  | 2 | 3 | 0.1528% | 1963 | Continuity Fail | 3
[8 empty] | 1 |  |  | 1 | 0.05094% | 1963 | Damaged Harness | 1
[8 empty] | 2 |  | 3 | 5 | 0.2547% | 1963 | Incorrect Weld | 5
[8 empty] | 4 |  |  | 4 | 0.2038% | 1963 | Missing Front/Back Stop Dot | 4
[11 empty] | 0 | 0% | 1963 | Desired Pattern Could Not Be Found | 0
[8 empty] | 2 |  |  | 2 | 0.1019% | 1963 | Fuse Switch | 2
[8 empty] | 3 |  |  | 3 | 0.1528% | 1963 | Exposed Wire | 3
[11 empty] | 0 | 0% | 1963 | Cosmetic Defect | 0
[8 empty] | 1 |  |  | 1 | 0.05094% | 1963 | Program Fail | 1
[8 empty] | 1 |  |  | 1 | 0.05094% | Build QTY | Pinch Cable | 1
[8 empty] | 4 |  |  | 4 | 0.2038% | 1963 | Damaged Component | 4 | 3ea Damaged Jaw(Dynamic) : 1ea Damaged Jaw (Static)
[14 empty] | Reworkables | 0
[9 empty] | 7 |  | 0 | 0% | 1963 | Blade Link Pop Out | 0
[9 empty] | 3 |  | 0 | 0% | 1963 | Scratched Logo | 0
[11 empty] | 0 | 0% | 1963 | Cosmetic Defect | 0
[9 empty] | 19 |  | 0 | 0% | 1963 | Missing Component | 0 | 15ea No Grease
[9 empty] | 3 |  | 0 | 0% | 1963 | Particulate Matter | 0
[9 empty] | 40 |  | 0 | 0% | 1963 | Misassembled Component | 0 | 11ea Excess Grease
[9 empty] | 10 |  | 0 | 0% | 1963 | Wire Popped Out | 0
[9 empty] | 3 |  | 0 | 0% | 1963 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[8 empty] | 1 |  |  | 0 | 0% | Build QTY | Excess Grivory | 0
|  |  | . | [7 empty] | 0 | 0% | Build QTY | Damaged omponent (Static) | 0
[11 empty] | 0 | 0% | 1963 | Cosmetic Defect | 0 | Scratched Logo :
[8 empty] | 4 |  |  | 4 | 0.2038% | Build QTY | Exposed Wire (Dynamic) | 4 | ea ARM COVER, STATIC
[8 empty] | 1 |  |  | 0 | 0% | 1963 | Exposed Wire (Static) | 0 | ea ARM HOUSING, STATIC
[8 empty] | 7 |  |  | 7 | 0.3566% | 1963 | Exposed Wire | 7
[8 empty] | 2 |  |  | 2 | 0.1019% | 1963 | Damaged Component | 2
[8 empty] | 5 |  |  | 5 | 0.2547% | 1963 | Front Stop not Touching | 5
[8 empty] | 6 |  |  | 6 | 0.3057% | 1963 | Blade Guide Does Not Go Through | 6
[8 empty] | 1 |  |  | 1 | 0.05094% | Build QTY | Misassembled Component | 1
[8 empty] | 4 |  |  | 4 | 0.2038% | 1963 | Incomplete / Incorrect Weld | 4
[8 empty] | 9 |  |  | 9 | 0.4585% | 1963 | Dark Spot Weld | 9
[11 empty] | 0 | 0% | 1963 | Incomplete / Incorrect Weld (Static) | 0
[11 empty] | 0 | 0% | 1963 | Missing Front/Back Stop Dot (Static) | 0
[11 empty] | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[8 empty] | 8 |  |  | 0 | 0% | 1963 | Cosmetic Defect | 0 | 6ea Missing Front Stop Dot
[7 empty] | Total | 164 | 85 | 22 | 176 | 8.966%
S/O | Model | S/O QTY | Build QTY | Yield | Yield % | Rework % | Date | Production Scrap | Production Rework | Quality Scrap | Total | Scrap % | Build Qty | Description | Total | Comments
1489769 | EB230 | 1920 | 1972 | 1770 | 89.76% | 4.108% | 2023-06-26 | [6 empty] | Top Level Non-Conformance |  | **Production Shop Order **
[8 empty] | 7 |  |  | 7 | 0.355% | 1972 | 0.003 Jaw Gap Fail | 7
[8 empty] | 4 |  |  | 4 | 0.2028% | 1972 | 0.007 Jaw Gap Fail | 4
[11 empty] | 0 | 0% | 1972 | High Jaw Force | 0
[11 empty] | 0 | 0% | 1972 | Low Jaw Force | 0
[8 empty] | 98 |  | 8 | 106 | 5.375% | 1972 | Stuck Blade | 106
[11 empty] | 0 | 0% | 1972 | Electrode Isolation | 0
[11 empty] | 0 | 0% | 1972 | Focus Error | 0
[8 empty] | 2 |  |  | 2 | 0.1014% | 1972 | Jaw Isolation | 2
[8 empty] | 15 |  | 1 | 16 | 0.8114% | 1972 | Continuity Fail | 16
[8 empty] | 1 |  |  | 1 | 0.05071% | 1972 | Damaged Harness | 1
[11 empty] | 0 | 0% | 1972 | Incorrect Weld | 0
[8 empty] | 1 |  |  | 1 | 0.05071% | 1972 | Missing Front/Back Stop Dot | 1
[11 empty] | 0 | 0% | 1972 | Desired Pattern Could Not Be Found | 0
[11 empty] | 0 | 0% | 1972 | Fuse Switch | 0
[8 empty] | 1 |  |  | 1 | 0.05071% | 1972 | Short Wire | 1
[8 empty] | 2 |  |  | 2 | 0.1014% | 1972 | Cosmetic Defect (Jaw) | 2
[11 empty] | 0 | 0% | 1972 | Program Fail | 0
[8 empty] | 3 |  | 6 | 9 | 0.4564% | Build QTY | Discolored Weld | 9
[8 empty] | 1 |  |  | 1 | 0.05071% | 1972 | Ben Jaw | 1
[14 empty] | Reworkables | 0
[9 empty] | 12 |  | 0 | 0% | 1972 | Blade Link Pop Out | 0
[9 empty] | 12 |  | 0 | 0% | 1972 | Scratched Logo | 0
[11 empty] | 0 | 0% | 1972 | Cosmetic Defect | 0
[9 empty] | 14 |  | 0 | 0% | 1972 | Missing Component | 0 | 12ea No Grease
[9 empty] | 2 |  | 0 | 0% | 1972 | Particulate Matter | 0
[9 empty] | 28 |  | 0 | 0% | 1972 | Misassembled Component | 0
[9 empty] | 13 |  | 0 | 0% | 1972 | Wire Popped Out | 0
[11 empty] | 0 | 0% | 1972 | Fuse Button | 0
[14 empty] | Sub-assembly Level Scrap | 0
[11 empty] | 0 | 0% | Build QTY | Excess Grivory | 0
|  |  | . |  |  |  |  | 1 |  |  | 0 | 0% | Build QTY | Damaged Component (Dynamic) | 0
[11 empty] | 0 | 0% | 1972 | Cosmetic Defect | 0 | Scratched Logo :
[8 empty] | 5 |  |  | 0 | 0% | Build QTY | Exposed Wire (Dynamic) | 0 | 31ea ARM COVER, STATIC
[8 empty] | 2 |  |  | 0 | 0% | 1972 | Exposed Wire (Static) | 0 | 30ea ARM HOUSING, STATIC
[8 empty] | 1 |  |  | 1 | 0.05071% | 1972 | Exposed Wire | 1
[8 empty] | 4 |  |  | 4 | 0.2028% | 1972 | Damaged Component | 4
[8 empty] | 17 |  |  | 17 | 0.8621% | 1972 | Front Stop not Touching | 17
[8 empty] | 11 |  |  | 11 | 0.5578% | 1972 | Blade Guide Does Not Go Through | 11
[8 empty] | 1 |  |  | 1 | 0.05071% | Build QTY | Misassembled Component | 1
[8 empty] | 2 |  |  | 2 | 0.1014% | 1972 | Incomplete / Incorrect Weld | 2
[8 empty] | 11 |  |  | 11 | 0.5578% | 1972 | Discolored Weld | 11
[11 empty] | 0 | 0% | 1972 | Incomplete / Incorrect Weld (Static) | 0 | 4ea Discolored Shaft
[8 empty] | 3 |  |  | 3 | 0.1521% | 1972 | Missing Front/Back Stop Dot | 3
[11 empty] | 0 | 0% | 588 | Excess Grivory (Dynamic) | 0
[8 empty] | 1 |  |  | 1 | 0.05071% | 1972 | Cosmetic Defect | 1
[7 empty] | Total | 194 | 81 | 15 | 201 | 10.19%

## EB040-EB240 Graphs
EB040/EB240 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 182
[5 empty] | Rough Jaw Actuation | 16
[5 empty] | High Jaw Force | 15
[5 empty] | Damaged Harness | 10
[5 empty] | 0.006 Jaw Gap Fail | 5
[5 empty] | Jaw Short | 5
[5 empty] | Front Stop Not Touching | 3
[5 empty] | Continuity Fail | 2
[5 empty] | Blade does not pass front stop | 1
[5 empty] | Fuse Switch | 1
[5 empty] | Incorrect/Missing Weld | 1
[5 empty] | Defective Component | 0
[5 empty] | Misassembled | 0
[5 empty] | Low Jaw Force | 0
[5 empty] | Rough Rotation Shaft Rotation | 0
[5 empty] | 0.003 Jaw Gap Fail | 0
Shop Order's Yield |  |  |  |  | Blade lever spring pop out | 0
Shop Order | Build QTY | Yield | Yield % | Date | Cracked Handles | 0
1493482 | 647 | 569 | 87.94% | 2023-06-07 | Trigger Lock Failure | 0
1492700 | 660 | 569 | 86.21% | 2023-06-09 | Gamma Audit | 0
1493483 | 637 | 565 | 88.7% | 2023-06-15
1492701 | 671 | 569 | 84.8% | 2023-06-19
1494748 | 646 | 570 | 88.24% | 2023-06-22
1493484 | 648 | 557 | 85.96% | 2023-06-27
Average Yield: |  |  | 86.97%

## EB040-EB240
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1473007 | EB240 | 576 | 606 | 568 | 93.73% | 2023-04-04 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 7 | [12 empty] | 0 | 0% | 568 | Rough Rotation Shaft Rotation | 0
[7 empty] | 4 | [13 empty] | 4 | 0.6601% | 568 | Rough Jaw Actuation | 4
[7 empty] | 21 | [13 empty] | 21 | 3.465% | 568 | Stuck Blade | 21
[21 empty] | 0 | 0% | 568 | Blade does not pass front stop | 0
[8 empty] | 2 | 1 | [11 empty] | 1 | 0.165% | 568 | 0.003 Jaw Gap Fail | 1
[20 empty] | 1 | 1 | 0.165% | 568 | 0.006 Jaw Gap Fail | 1
[8 empty] | 39 | 3 | [11 empty] | 3 | 0.495% | 568 | High Jaw Force | 3
[21 empty] | 0 | 0% | 568 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 568 | Front Stop Not Touching | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 568 | Jaw Short | 0
[21 empty] | 0 | 0% | 568 | Fuse Switch | 0
[20 empty] | 1 | 1 | 0.165% | 568 | Continuity Fail | 1
[21 empty] | 0 | 0% | 568 | Cosmetic Defect | 0
[21 empty] | 0 | 0% | 568 | Incorrect Blade Orientation | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 568 | Misassembled | 0
[8 empty] | 6 | [12 empty] | 0 | 0% | 568 | Missing Component | 0 | 5ea No Grease
[8 empty] | 2 | [12 empty] | 0 | 0% | 568 | Failure To Unblock Trigger | 0
[9 empty] | 2 | [11 empty] | 2 | 0.33% | 568 | Damaged Harness | 2
[21 empty] | 0 | 0% | 568 | Scratched Logo | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 568 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[21 empty] | 0 | 0% | 568 | Damaged Frame | 0
[21 empty] | 0 | 0% | 568 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 568 | Blade guide assembled incorrectly | 0
[21 empty] | 0 | 0% | 568 | Damaged actuation tube from nut setting | 0
[21 empty] | 0 | 0% | 568 | Incorrect Weld | 0
[7 empty] | 2 | [13 empty] | 2 | 0.33% | 568 | Exposed Wire | 2
[21 empty] | 0 | 0% | 568 | Frames Weld Incorrectly | 0
[21 empty] | 0 | 0% | 568 | Jaws Do Not Open | 0 | 3ea Defective Tube
[21 empty] | 0 | 0% | 568 | Misaligned Hub | 0 | 4ea Defective Upper Jaw
[7 empty] | 1 | [13 empty] | 1 | 0.165% | 568 | Broken Nut Thread | 1 | 18ea Damaged Hub
[21 empty] | 0 | 0% | 568 | Front Stop Not Touching | 0
[7 empty] | 2 | [13 empty] | 2 | 0.33% | 568 | Misassembled Frames | 2
[21 empty] | 0 | 0% | 568 | Damaged Component | 0
[6 empty] | Total | 30 | 59 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 38 | 6.271% | 568
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1475189 | EB240 | 576 | 657 | 560 | 85.24% | 2023-04-18 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 21 |  | 1 | [10 empty] | 0 | 0% | 560 | Rough Rotation Shaft Rotation | 0
[7 empty] | 9 |  | 7 |  | 2 | [8 empty] | 1 | 19 | 2.892% | 560 | Rough Jaw Actuation | 19
[7 empty] | 49 |  | 3 |  | 4 | [9 empty] | 56 | 8.524% | 560 | Stuck Blade | 56
[21 empty] | 0 | 0% | 560 | Blade does not pass front stop | 0
[8 empty] | 2 | [12 empty] | 0 | 0% | 560 | 0.003 Jaw Gap Fail | 0
[8 empty] | 1 | 2 | [11 empty] | 2 | 0.3044% | 560 | 0.006 Jaw Gap Fail | 2
[8 empty] | 1 |  | 3 | 4 | [9 empty] | 4 | 0.6088% | 560 | High Jaw Force | 4
[21 empty] | 0 | 0% | 560 | Low Jaw Force | 0
[8 empty] | 1 | 1 | [11 empty] | 1 | 0.1522% | 560 | Front Stop Not Touching | 1
[8 empty] | 4 | 3 | 1 | 1 | [9 empty] | 4 | 0.6088% | 560 | Jaw Short | 4
[21 empty] | 0 | 0% | 560 | Fuse Switch | 0
[8 empty] | 4 | 2 | 1 | [10 empty] | 2 | 0.3044% | 560 | Continuity Fail | 2
[21 empty] | 0 | 0% | 560 | Cosmetic Defect | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1522% | 560 | Incorrect Blade Orientation | 1
[8 empty] | 15 |  | 3 | [10 empty] | 0 | 0% | 560 | Misassembled | 0
[8 empty] | 6 | [12 empty] | 0 | 0% | 560 | Missing Component | 0 | 1ea No Grease
[21 empty] | 0 | 0% | 560 | Failure To Unblock Trigger | 0
[11 empty] | 1 | [9 empty] | 1 | 0.1522% | 560 | Damaged Harness | 1
[8 empty] | 3 | [11 empty] | 3 | 3 | 0.4566% | 560 | Scratched Logo | 3
[8 empty] | 1 | [12 empty] | 0 | 0% | 560 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[21 empty] | 0 | 0% | 560 | Damaged Frame | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1522% | 560 | Incorrect Hub Press | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1522% | 560 | Blade guide assembled incorrectly | 1
[21 empty] | 0 | 0% | 560 | Damaged actuation tube from nut setting | 0
[21 empty] | 0 | 0% | 560 | Incorrect Weld | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1522% | 560 | Exposed Wire | 1
[21 empty] | 0 | 0% | 560 | Frames Weld Incorrectly | 0
[21 empty] | 0 | 0% | 560 | Jaws Do Not Open | 0 | 14ea Defective Tube
[7 empty] | 1 | [13 empty] | 1 | 0.1522% | 560 | Misaligned Hub | 1 | 3ea Defective Upper Jaw
[21 empty] | 0 | 0% | 560 | Broken Nut Thread | 0 | 28ea Damaged Hub
[21 empty] | 0 | 0% | 560 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 560 | Misassembled Frames | 0
[21 empty] | 0 | 0% | 560 | Damaged Component | 0
[6 empty] | Total | 63 | 59 | 18 | 9 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 97 | 14.76% | 560
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1486222 | EB240 | 576 | 627 | 554 | 88.36% | 2023-04-24 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 7 |  | 2 | 1 | [9 empty] | 1 | 0.1595% | 554 | Rough Rotation Shaft Rotation | 1
[7 empty] | 7 |  | 1 | [11 empty] | 8 | 1.276% | 554 | Rough Jaw Actuation | 8
[7 empty] | 25 |  | 2 |  | 3 | [9 empty] | 30 | 4.785% | 554 | Stuck Blade | 30
[21 empty] | 0 | 0% | 554 | Blade does not pass front stop | 0
[8 empty] | 10 | 5 | 1 | 1 | [8 empty] | 6 | 12 | 1.914% | 554 | 0.003 Jaw Gap Fail | 12
[8 empty] | 1 | 2 | [11 empty] | 2 | 0.319% | 554 | 0.006 Jaw Gap Fail | 2
[8 empty] | 1 |  | 4 | 4 | [9 empty] | 4 | 0.638% | 554 | High Jaw Force | 4
[21 empty] | 0 | 0% | 554 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 554 | Front Stop Not Touching | 0
[10 empty] | 1 | 1 | [9 empty] | 1 | 0.1595% | 554 | Jaw Short | 1
[21 empty] | 0 | 0% | 554 | Fuse Switch | 0
[8 empty] | 1 | 1 | 1 | 1 | [9 empty] | 2 | 0.319% | 554 | Continuity Fail | 2
[21 empty] | 0 | 0% | 554 | Cosmetic Defect | 0
[21 empty] | 0 | 0% | 554 | Incorrect Blade Orientation | 0
[8 empty] | 25 | [12 empty] | 0 | 0% | 554 | Misassembled | 0
[8 empty] | 3 |  | 1 | 1 | [9 empty] | 1 | 0.1595% | 554 | Missing Component | 1 | 2ea No Grease
[8 empty] | 1 | [12 empty] | 0 | 0% | 554 | Failure To Unblock Trigger | 0
[11 empty] | 4 | [9 empty] | 4 | 0.638% | 554 | Damaged Harness | 4
[8 empty] | 1 | [11 empty] | 1 | 1 | 0.1595% | 554 | Scratched Logo | 1
[21 empty] | 0 | 0% | 554 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[21 empty] | 0 | 0% | 554 | Damaged Frame | 0
[21 empty] | 0 | 0% | 554 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 554 | Blade guide assembled incorrectly | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1595% | 554 | Damaged actuation tube from nut setting | 1
[21 empty] | 0 | 0% | 554 | Incorrect Weld | 0
[7 empty] | 4 | [13 empty] | 4 | 0.638% | 554 | Exposed Wire | 4
[21 empty] | 0 | 0% | 554 | Frames Weld Incorrectly | 0
[21 empty] | 0 | 0% | 554 | Jaws Do Not Open | 0 | 2ea Defective Tube
[7 empty] | 2 | [13 empty] | 2 | 0.319% | 554 | Misaligned Hub | 2 | 2ea Defective Upper Jaw
[21 empty] | 0 | 0% | 554 | Broken Nut Thread | 0
[21 empty] | 0 | 0% | 554 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 554 | Misassembled Frames | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1595% | 554 | Damaged Component | 1
[6 empty] | Total | 40 | 50 | 11 | 10 | 16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 74 | 11.8% | 554
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1486223 | EB240 | 576 | 626 | 558 | 89.14% | 2023-04-28 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 7 | [12 empty] | 0 | 0% | 558 | Rough Rotation Shaft Rotation | 0
[7 empty] | 3 | [13 empty] | 3 | 0.4792% | 558 | Rough Jaw Actuation | 3
[7 empty] | 43 |  | 5 | [10 empty] | 4 | 52 | 8.307% | 558 | Stuck Blade | 52
[21 empty] | 0 | 0% | 558 | Blade does not pass front stop | 0
[8 empty] | 5 | 3 | [11 empty] | 3 | 0.4792% | 558 | 0.003 Jaw Gap Fail | 3
[8 empty] | 1 | 1 | [11 empty] | 1 | 0.1597% | 558 | 0.006 Jaw Gap Fail | 1
[21 empty] | 0 | 0% | 558 | High Jaw Force | 0
[21 empty] | 0 | 0% | 558 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 558 | Front Stop Not Touching | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 558 | Jaw Short | 0
[21 empty] | 0 | 0% | 558 | Fuse Switch | 0
[8 empty] | 1 | 1 | [11 empty] | 1 | 0.1597% | 558 | Continuity Fail | 1
[21 empty] | 0 | 0% | 558 | Cosmetic Defect | 0
[21 empty] | 0 | 0% | 558 | Incorrect Blade Orientation | 0
[8 empty] | 8 | [12 empty] | 0 | 0% | 558 | Misassembled | 0
[8 empty] | 5 | [12 empty] | 0 | 0% | 558 | Missing Component | 0 | 3ea No Grease
[8 empty] | 4 | 1 | [11 empty] | 1 | 0.1597% | 558 | Failure To Unblock Trigger | 1
[9 empty] | 3 | [11 empty] | 3 | 0.4792% | 558 | Damaged Harness | 3
[8 empty] | 1 | [12 empty] | 0 | 0% | 558 | Scratched Logo | 0
[21 empty] | 0 | 0% | 558 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[21 empty] | 0 | 0% | 558 | Damaged Frame | 0
[21 empty] | 0 | 0% | 558 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 558 | Blade guide assembled incorrectly | 0
[21 empty] | 0 | 0% | 558 | Damaged actuation tube from nut setting | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1597% | 558 | Incorrect Weld | 1
[21 empty] | 0 | 0% | 558 | Misassembled | 0
[21 empty] | 0 | 0% | 558 | Frames Weld Incorrectly | 0
[21 empty] | 0 | 0% | 558 | Jaws Do Not Open | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1597% | 558 | Misaligned Hub | 1
[21 empty] | 0 | 0% | 558 | Broken Nut Thread | 0
[21 empty] | 0 | 0% | 558 | Front Stop Not Touching | 0
[7 empty] | 2 | [13 empty] | 2 | 0.3195% | 558 | Misassembled Frames | 2
[21 empty] | 0 | 0% | 558 | Damaged Component | 0
[6 empty] | Total | 50 | 33 | 14 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 68 | 10.86% | 558
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1486767 | EB240 | 576 | 608 | 564 | 92.76% | 2023-05-08 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 11 |  | 4 | [10 empty] | 0 | 0% | 564 | Rough Rotation Shaft Rotation | 0
[21 empty] | 0 | 0% | 564 | Rough Jaw Actuation | 0
[7 empty] | 23 | [13 empty] | 23 | 3.783% | 564 | Stuck Blade | 23
[21 empty] | 0 | 0% | 564 | Blade does not pass front stop | 0
[10 empty] | 1 | [10 empty] | 0 | 0% | 564 | 0.003 Jaw Gap Fail | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 564 | 0.006 Jaw Gap Fail | 0
[8 empty] | 5 |  | 1 | [10 empty] | 0 | 0% | 564 | High Jaw Force | 0
[21 empty] | 0 | 0% | 564 | Low Jaw Force | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 564 | Front Stop Not Touching | 0
[20 empty] | 1 | 1 | 0.1645% | 564 | Jaw Short | 1
[8 empty] | 1 | [12 empty] | 0 | 0% | 564 | Fuse Switch | 0
[8 empty] | 4 | [12 empty] | 0 | 0% | 564 | Continuity Fail | 0
[21 empty] | 0 | 0% | 564 | Cosmetic Defect | 0
[21 empty] | 0 | 0% | 564 | Incorrect Blade Orientation | 0
[8 empty] | 9 | [12 empty] | 0 | 0% | 564 | Misassembled | 0
[8 empty] | 13 |  | 1 | [10 empty] | 0 | 0% | 564 | Missing Component | 0 | 3ea No Grease
[8 empty] | 14 |  | 2 | [10 empty] | 0 | 0% | 564 | Failure To Unblock Trigger | 0
[11 empty] | 7 | [7 empty] | 2 |  | 9 | 1.48% | 564 | Damaged Harness | 9 | 2ea Lost Serial Number
[19 empty] | 1 |  | 1 | 0.1645% | 564 | Scratched Logo | 1
[19 empty] | 1 |  | 1 | 0.1645% | 564 | Particulate Matter | 1
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 3 | [13 empty] | 3 | 0.4934% | 564 | Exposed Wire | 3
[7 empty] | 1 | [13 empty] | 1 | 0.1645% | 564 | Incorrect Hub Press | 1
[21 empty] | 0 | 0% | 564 | Blade guide assembled incorrectly | 0
[21 empty] | 0 | 0% | 564 | Damaged actuation tube from nut setting | 0
[7 empty] | 3 | [13 empty] | 3 | 0.4934% | 564 | Incorrect Weld | 3
[7 empty] | 2 | [13 empty] | 2 | 0.3289% | 564 | Misassembled | 2
[21 empty] | 0 | 0% | 564 | Frames Weld Incorrectly | 0
[21 empty] | 0 | 0% | 564 | Jaws Do Not Open | 0
[21 empty] | 0 | 0% | 564 | Misaligned Hub | 0
[21 empty] | 0 | 0% | 564 | Broken Nut Thread | 0
[21 empty] | 0 | 0% | 564 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 564 | Misassembled Frames | 0
[21 empty] | 0 | 0% | 564 | Damaged Component | 0
[6 empty] | Total | 32 | 59 | 0 | 9 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 1 | 40 | 6.579% | 564
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1488017 | EB240 | 576 | 649 | 564 | 86.9% | 2023-05-12 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 17 | [12 empty] | 0 | 0% | 564 | Rough Rotation Shaft Rotation | 0
[7 empty] | 3 | [13 empty] | 3 | 0.4622% | 564 | Rough Jaw Actuation | 3
[7 empty] | 46 |  | 12 |  | 1 | [9 empty] | 59 | 9.091% | 564 | Stuck Blade | 59
[21 empty] | 0 | 0% | 564 | Blade does not pass front stop | 0
[9 empty] | 1 |  | 2 | [9 empty] | 3 | 0.4622% | 564 | 0.003 Jaw Gap Fail | 3
[7 empty] | 1 | 1 | 1 | 1 | 2 | [9 empty] | 4 | 0.6163% | 564 | 0.006 Jaw Gap Fail | 4
[8 empty] | 2 |  | 2 | 1 | [9 empty] | 1 | 0.1541% | 564 | High Jaw Force | 1
[8 empty] | 2 | [12 empty] | 0 | 0% | 564 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 564 | Front Stop Not Touching | 0
[8 empty] | 3 |  | 1 | [10 empty] | 0 | 0% | 564 | Jaw Short | 0
[21 empty] | 0 | 0% | 564 | Fuse Switch | 0
[8 empty] | 2 |  | 2 | 2 | [9 empty] | 2 | 0.3082% | 564 | Continuity Fail | 2
[11 empty] | 1 | [9 empty] | 1 | 0.1541% | 564 | Failure To Unlock Trigger | 1
[21 empty] | 0 | 0% | 564 | Incorrect Blade Orientation | 0
[8 empty] | 23 |  | 2 | [10 empty] | 0 | 0% | 564 | Misassembled | 0
[8 empty] | 11 | [12 empty] | 0 | 0% | 564 | Missing Component | 0 | 8ea No Grease
[8 empty] | 19 | [12 empty] | 0 | 0% | 564 | Failure To lock Trigger | 0
[11 empty] | 3 | [9 empty] | 3 | 0.4622% | 564 | Damaged Harness | 3
[8 empty] | 1 | [12 empty] | 0 | 0% | 564 | Scratched Logo | 0
[21 empty] | 0 | 0% | 564 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1541% | 564 | Exposed Wire | 1
[21 empty] | 0 | 0% | 564 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 564 | Blade guide assembled incorrectly | 0
[21 empty] | 0 | 0% | 564 | Damaged actuation tube from nut setting | 0 | 4ea Exposed Wire (Upper Jaw)
[7 empty] | 1 | [13 empty] | 1 | 0.1541% | 564 | Incorrect Weld | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1541% | 564 | Misassembled | 1
[21 empty] | 0 | 0% | 564 | Frames Weld Incorrectly | 0
[7 empty] | 2 | [13 empty] | 2 | 0.3082% | 564 | Jaws Do Not Open | 2
[7 empty] | 3 | [13 empty] | 3 | 0.4622% | 564 | Misaligned Hub | 3
[21 empty] | 0 | 0% | 564 | Broken Nut Thread | 0
[21 empty] | 0 | 0% | 564 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 564 | Misassembled Frames | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1541% | 564 | Damaged Component | 1 | Short Wire
[6 empty] | Total | 59 | 81 | 14 | 8 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 85 | 13.1% | 564
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1488944 | EB240 | 576 | 649 | 557 | 85.82% | 2023-05-16 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 21 | [12 empty] | 0 | 0% | 557 | Rough Rotation Shaft Rotation | 0
[7 empty] | 7 | [13 empty] | 7 | 1.079% | 557 | Rough Jaw Actuation | 7
[7 empty] | 42 |  | 2 | [11 empty] | 44 | 6.78% | 557 | Stuck Blade | 44
[21 empty] | 0 | 0% | 557 | Blade does not pass front stop | 0
[8 empty] | 3 | 6 | [11 empty] | 6 | 0.9245% | 557 | 0.003 Jaw Gap Fail | 6
[8 empty] | 4 | 4 | [11 empty] | 4 | 0.6163% | 557 | 0.006 Jaw Gap Fail | 4
[8 empty] | 14 | 3 | [11 empty] | 3 | 0.4622% | 557 | High Jaw Force | 3
[8 empty] | 1 | [12 empty] | 0 | 0% | 557 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 557 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 557 | Jaw Short | 0
[21 empty] | 0 | 0% | 557 | Fuse Switch | 0
[8 empty] | 3 | 1 | [11 empty] | 1 | 0.1541% | 557 | Continuity Fail | 1
[8 empty] | 4 | [12 empty] | 0 | 0% | 557 | Failure To Unlock Trigger | 0
[21 empty] | 0 | 0% | 557 | Incorrect Blade Orientation | 0
[8 empty] | 12 | [12 empty] | 0 | 0% | 557 | Misassembled | 0
[8 empty] | 4 | [12 empty] | 0 | 0% | 557 | Missing Component | 0 | 1ea No Grease
[19 empty] | 1 |  | 1 | 0.1541% | 557 | Cosmetic Defect | 1
[9 empty] | 1 | [11 empty] | 1 | 0.1541% | 557 | Damaged Harness | 1
[8 empty] | 1 | [11 empty] | 1 | 1 | 0.1541% | 557 | Scratched Logo | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1541% | 557 | Damaged Shaft | 1
[24 empty] | Shaft Level Scrap | 0
[21 empty] | 0 | 0% | 557 | Exposed Wire | 0
[21 empty] | 0 | 0% | 557 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 557 | Blade guide assembled incorrectly | 0
[21 empty] | 0 | 0% | 557 | Damaged actuation tube from nut setting | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1541% | 557 | Incorrect Weld | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1541% | 557 | Misassembled | 1
[21 empty] | 0 | 0% | 557 | Frames Weld Incorrectly | 0
[7 empty] | 2 | [13 empty] | 2 | 0.3082% | 557 | Jaws Do Not Open | 2 | 9ea Damaged Tube
[7 empty] | 2 | [13 empty] | 2 | 0.3082% | 557 | Misaligned Hub | 2 | 1ea Exposed Wire Upper Jaw
[7 empty] | 8 | [13 empty] | 8 | 1.233% | 557 | Rough Jaw | 8
[21 empty] | 0 | 0% | 557 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 557 | Misassembled Frames | 0
[7 empty] | 9 | [13 empty] | 9 | 1.387% | 557 | Stuck Blade | 9
[6 empty] | Total | 73 | 67 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 91 | 14.02% | 557
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1488014 | EB240 | 576 | 734 | 560 | 76.29% | 2023-05-20 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 9 | [12 empty] | 0 | 0% | 560 | Rough Rotation Shaft Rotation | 0
[7 empty] | 13 |  | 6 | [11 empty] | 19 | 2.589% | 560 | Rough Jaw Actuation | 19
[7 empty] | 94 |  | 7 |  | 7 | [9 empty] | 108 | 14.71% | 560 | Stuck Blade | 108
[21 empty] | 0 | 0% | 560 | Blade does not pass front stop | 0
[8 empty] | 1 | 5 | 2 | 1 | [9 empty] | 6 | 0.8174% | 560 | 0.003 Jaw Gap Fail | 6
[8 empty] | 2 | 2 | 1 | [10 empty] | 2 | 0.2725% | 560 | 0.006 Jaw Gap Fail | 2
[8 empty] | 12 | 4 | 4 | 5 | [9 empty] | 9 | 1.226% | 560 | High Jaw Force | 9
[21 empty] | 0 | 0% | 560 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 560 | Front Stop Not Touching | 0
[8 empty] | 4 | 1 | [11 empty] | 1 | 0.1362% | 560 | Jaw Short | 1
[21 empty] | 0 | 0% | 560 | Fuse Switch | 0
[21 empty] | 0 | 0% | 560 | Continuity Fail | 0
[8 empty] | 6 | [12 empty] | 0 | 0% | 560 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 560 | Incorrect Blade Orientation | 0
[8 empty] | 18 | [12 empty] | 0 | 0% | 560 | Misassembled | 0
[8 empty] | 8 |  | 2 | [10 empty] | 0 | 0% | 560 | Missing Component | 0 | 2ea No Grease
[21 empty] | 0 | 0% | 560 | Failure To lock Trigger | 0
[11 empty] | 3 | [9 empty] | 3 | 0.4087% | 560 | Damaged Harness | 3
[21 empty] | 0 | 0% | 560 | Scratched Logo | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 560 | Stuck Button | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 3 | [13 empty] | 3 | 0.4087% | 560 | Exposed Wire | 3
[21 empty] | 0 | 0% | 560 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 560 | Blade guide assembled incorrectly | 0
[7 empty] | 2 | [13 empty] | 2 | 0.2725% | 560 | Damaged actuation tube from nut setting | 2
[21 empty] | 0 | 0% | 560 | Incorrect Weld | 0
[7 empty] | 8 | [13 empty] | 8 | 1.09% | 560 | Misassembled | 8
[7 empty] | 1 | [13 empty] | 1 | 0.1362% | 560 | Frames Weld Incorrectly | 1
[7 empty] | 2 | [13 empty] | 2 | 0.2725% | 560 | Jaws Do Not Open | 2 | 3ea Damaged Tube
[7 empty] | 1 | [13 empty] | 1 | 0.1362% | 560 | Misaligned Hub | 1
[7 empty] | 9 | [13 empty] | 9 | 1.226% | 560 | Rough Jaw | 9
[21 empty] | 0 | 0% | 560 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 560 | Misassembled Frames | 0
[21 empty] | 0 | 0% | 560 | Stuck Blade | 0
[6 empty] | Total | 133 | 61 | 25 | 9 | 16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 174 | 23.71% | 560
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1491100 | EB240 | 576 | 674 | 565 | 83.83% | 2023-05-23 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 6 |  | 1 | [10 empty] | 0 | 0% | 565 | Rough Rotation Shaft Rotation | 0
[7 empty] | 10 |  | 6 |  | 1 | [9 empty] | 17 | 2.522% | 565 | Rough Jaw Actuation | 17
[7 empty] | 44 |  | 12 |  | 3 | [9 empty] | 59 | 8.754% | 565 | Stuck Blade | 59
[21 empty] | 0 | 0% | 565 | Blade does not pass front stop | 0
[8 empty] | 8 | 4 | 1 | 1 | [8 empty] | 1 | 6 | 0.8902% | 565 | 0.003 Jaw Gap Fail | 6
[21 empty] | 0 | 0% | 565 | 0.006 Jaw Gap Fail | 0
[8 empty] | 4 | 1 | 1 | [10 empty] | 1 | 0.1484% | 565 | High Jaw Force | 1
[21 empty] | 0 | 0% | 565 | Low Jaw Force | 0
[8 empty] | 2 | 1 | [11 empty] | 1 | 0.1484% | 565 | Front Stop Not Touching | 1
[21 empty] | 0 | 0% | 565 | Jaw Short | 0
[21 empty] | 0 | 0% | 565 | Fuse Switch | 0
[8 empty] | 2 | 2 | [10 empty] | 1 | 3 | 0.4451% | 565 | Continuity Fail | 3
[8 empty] | 23 | 2 | [11 empty] | 2 | 0.2967% | 565 | Failure To Lock/Unlock Trigger | 2
[21 empty] | 0 | 0% | 565 | Incorrect Blade Orientation | 0
[8 empty] | 7 | [12 empty] | 0 | 0% | 565 | Misassembled | 0
[8 empty] | 6 | [12 empty] | 0 | 0% | 565 | Missing Component | 0 | 3ea No Grease
[21 empty] | 0 | 0% | 565 | Failure To lock Trigger | 0
[11 empty] | 2 | [7 empty] | 2 |  | 4 | 0.5935% | 565 | Damaged Harness | 4
[21 empty] | 0 | 0% | 565 | Scratched Logo | 0
[21 empty] | 0 | 0% | 565 | Stuck Button | 0
[24 empty] | Shaft Level Scrap | 0
[21 empty] | 0 | 0% | 565 | Exposed Wire | 0
[21 empty] | 0 | 0% | 565 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 565 | Blade guide assembled incorrectly | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1484% | 565 | Damaged actuation tube from nut setting | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1484% | 565 | Incorrect Weld | 1
[7 empty] | 5 | [13 empty] | 5 | 0.7418% | 565 | Misassembled | 5
[7 empty] | 1 | [13 empty] | 1 | 0.1484% | 565 | Frames Weld Incorrectly | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1484% | 565 | Damaged Thread | 1
[7 empty] | 2 | [13 empty] | 2 | 0.2967% | 565 | Misaligned Hub | 2
[7 empty] | 4 | [13 empty] | 4 | 0.5935% | 565 | Rough Jaw | 4
[21 empty] | 0 | 0% | 565 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 565 | Misassembled Frames | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1484% | 565 | Damaged Component | 1 | Crack Jaw
[6 empty] | Total | 70 | 58 | 28 | 3 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 107 | 15.88% | 565
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1491101 | EB240 | 576 | 660 | 559 | 84.7% | 2023-05-25 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 10 | [12 empty] | 0 | 0% | 559 | Rough Rotation Shaft Rotation | 0
[7 empty] | 10 |  | 3 | [11 empty] | 13 | 1.97% | 559 | Rough Jaw Actuation | 13
[7 empty] | 63 |  | 4 | [11 empty] | 67 | 10.15% | 559 | Stuck Blade | 67
[21 empty] | 0 | 0% | 559 | Blade does not pass front stop | 0
[8 empty] | 4 | 1 | [10 empty] | 1 | 2 | 0.303% | 559 | 0.003 Jaw Gap Fail | 2
[21 empty] | 0 | 0% | 559 | 0.006 Jaw Gap Fail | 0
[8 empty] | 14 | 1 | [11 empty] | 1 | 0.1515% | 559 | High Jaw Force | 1
[21 empty] | 0 | 0% | 559 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 559 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 559 | Jaw Short | 0
[21 empty] | 0 | 0% | 559 | Fuse Switch | 0
[8 empty] | 1 | 1 | [10 empty] | 1 | 2 | 0.303% | 559 | Continuity Fail | 2
[8 empty] | 13 | [12 empty] | 0 | 0% | 559 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 559 | Incorrect Blade Orientation | 0
[8 empty] | 10 | [12 empty] | 0 | 0% | 559 | Misassembled | 0
[8 empty] | 8 | [12 empty] | 0 | 0% | 559 | Missing Component | 0 | 3ea No Grease
[21 empty] | 0 | 0% | 559 | Failure To lock Trigger | 0
[9 empty] | 3 | [11 empty] | 3 | 0.4545% | 559 | Damaged Harness | 3
[8 empty] | 3 | [11 empty] | 1 | 1 | 0.1515% | 559 | Scratched Logo | 1
[19 empty] | 1 |  | 1 | 0.1515% | 559 | Engineering Sample | 1
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.303% | 559 | Exposed Wire | 2
[21 empty] | 0 | 0% | 559 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 559 | Blade guide assembled incorrectly | 0
[7 empty] | 2 | [13 empty] | 2 | 0.303% | 559 | Damaged actuation tube from nut setting | 2
[21 empty] | 0 | 0% | 559 | Incorrect Weld | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 559 | Broken Blade Pusher | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 559 | Frames Weld Incorrectly | 1
[21 empty] | 0 | 0% | 559 | Damaged Thread | 0
[21 empty] | 0 | 0% | 559 | Misaligned Hub | 0
[7 empty] | 2 | [13 empty] | 2 | 0.303% | 559 | Jaws Do Not Open | 2
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 559 | Front Stop Not Touching | 1
[7 empty] | 2 | [13 empty] | 2 | 0.303% | 559 | Misassembled | 2
[21 empty] | 0 | 0% | 559 | Damaged Component | 0
[6 empty] | Total | 84 | 63 | 13 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 3 | 100 | 15.15% | 559
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1491102 | EB240 | 576 | 666 | 567 | 85.14% | 2023-05-27 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 11 |  | 1 | [10 empty] | 0 | 0% | 567 | Rough Rotation Shaft Rotation | 0
[7 empty] | 11 |  | 4 | [11 empty] | 15 | 2.252% | 567 | Rough Jaw Actuation | 15
[7 empty] | 46 |  | 10 |  | 1 | [8 empty] | 1 | 58 | 8.709% | 567 | Stuck Blade | 58
[21 empty] | 0 | 0% | 567 | Blade does not pass front stop | 0
[8 empty] | 7 | 2 | [10 empty] | 1 | 3 | 0.4505% | 567 | 0.003 Jaw Gap Fail | 3
[8 empty] | 1 | 1 | [11 empty] | 1 | 0.1502% | 567 | 0.006 Jaw Gap Fail | 1
[8 empty] | 42 | 8 | 6 | 4 | [9 empty] | 12 | 1.802% | 567 | High Jaw Force | 12
[21 empty] | 0 | 0% | 567 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 567 | Front Stop Not Touching | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 567 | Jaw Short | 0
[21 empty] | 0 | 0% | 567 | Fuse Switch | 0
[21 empty] | 0 | 0% | 567 | Continuity Fail | 0
[8 empty] | 29 | [12 empty] | 0 | 0% | 567 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 567 | Incorrect Blade Orientation | 0
[8 empty] | 10 | [12 empty] | 0 | 0% | 567 | Misassembled | 0
[8 empty] | 9 | [12 empty] | 0 | 0% | 567 | Missing Component | 0 | 0ea No Grease
[21 empty] | 0 | 0% | 567 | Failure To lock Trigger | 0
[11 empty] | 2 | [9 empty] | 2 | 0.3003% | 567 | Damaged Harness | 2
[8 empty] | 1 | [12 empty] | 0 | 0% | 567 | Scratched Logo | 0
[21 empty] | 0 | 0% | 567 | Missing Harness | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.3003% | 567 | Exposed Wire | 2
[21 empty] | 0 | 0% | 567 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 567 | Blade guide assembled incorrectly | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1502% | 567 | Damaged actuation tube from nut setting | 1
[21 empty] | 0 | 0% | 567 | Incorrect Weld | 0
[21 empty] | 0 | 0% | 567 | Broken Blade Pusher | 0
[21 empty] | 0 | 0% | 567 | Frames Weld Incorrectly | 0
[21 empty] | 0 | 0% | 567 | Damaged Thread | 0 | 25ea Damaged Lower Jaw
[7 empty] | 1 | [13 empty] | 1 | 0.1502% | 567 | Misaligned Hub | 1 | 20ea Damaged Upper Jaw
[7 empty] | 3 | [13 empty] | 3 | 0.4505% | 567 | Jaws Do Not Open | 3
[21 empty] | 0 | 0% | 567 | Front Stop Not Touching | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1502% | 567 | Misassembled | 1
[21 empty] | 0 | 0% | 567 | Damaged Component | 0
[6 empty] | Total | 65 | 111 | 25 | 7 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 99 | 14.86% | 567
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1488015 | EB240 | 576 | 637 | 567 | 89.01% | 2023-06-01 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 12 | 1 |  | 1 | [9 empty] | 2 | 0.314% | 567 | Rough Rotation Shaft Rotation | 2
[7 empty] | 1 | [13 empty] | 1 | 0.157% | 567 | Rough Jaw Actuation | 1
[7 empty] | 37 |  | 3 | [11 empty] | 40 | 6.279% | 567 | Stuck Blade | 40
[21 empty] | 0 | 0% | 567 | Blade does not pass front stop | 0
[8 empty] | 8 | 2 | [10 empty] | 1 | 3 | 0.471% | 567 | 0.003 Jaw Gap Fail | 3
[8 empty] | 2 | 2 |  | 1 | [9 empty] | 3 | 0.471% | 567 | 0.006 Jaw Gap Fail | 3
[8 empty] | 6 |  | 3 | 3 | [9 empty] | 3 | 0.471% | 567 | High Jaw Force | 3
[21 empty] | 0 | 0% | 567 | Low Jaw Force | 0
[9 empty] | 1 | [11 empty] | 1 | 0.157% | 567 | Front Stop Not Touching | 1
[21 empty] | 0 | 0% | 567 | Jaw Short | 0
[21 empty] | 0 | 0% | 567 | Fuse Switch | 0
[21 empty] | 0 | 0% | 567 | Continuity Fail | 0
[8 empty] | 2 | 2 |  | 1 | [9 empty] | 3 | 0.471% | 567 | Failure To Lock/Unlock Trigger | 3
[21 empty] | 0 | 0% | 567 | Incorrect Blade Orientation | 0
[8 empty] | 11 | [12 empty] | 0 | 0% | 567 | Misassembled | 0
[8 empty] | 3 | [12 empty] | 0 | 0% | 567 | Missing Component | 0 | 2ea No Grease
[8 empty] | 1 | [12 empty] | 0 | 0% | 567 | Cosmetic Defect | 0
[11 empty] | 2 | [9 empty] | 2 | 0.314% | 567 | Damaged Harness | 2
[8 empty] | 2 | [12 empty] | 0 | 0% | 567 | Scratched Logo | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 567 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.314% | 567 | Exposed Wire | 2
[21 empty] | 0 | 0% | 567 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 567 | Blade guide assembled incorrectly | 0
[21 empty] | 0 | 0% | 567 | Damaged actuation tube from nut setting | 0
[7 empty] | 7 | [13 empty] | 7 | 1.099% | 567 | Incorrect Weld | 7
[7 empty] | 1 | [13 empty] | 1 | 0.157% | 567 | Stuck Blade | 1
[21 empty] | 0 | 0% | 567 | Frames Weld Incorrectly | 0 | 2ea Damaged Tube
[21 empty] | 0 | 0% | 567 | Damaged Thread | 0 | 60ea Damaged Lower Jaw
[7 empty] | 2 | [13 empty] | 2 | 0.314% | 567 | Misaligned Hub | 2 | 24ea Damaged Upper Jaw
[21 empty] | 0 | 0% | 567 | Jaws Do Not Open | 0
[21 empty] | 0 | 0% | 567 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 567 | Misassembled | 0
[21 empty] | 0 | 0% | 567 | Damaged Component | 0
[6 empty] | Total | 50 | 48 | 11 | 3 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 70 | 10.99% | 567
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1491103 | EB240 | 576 | 672 | 567 | 84.38% | 2023-06-03 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 14 |  | 1 | [10 empty] | 0 | 0% | 567 | Rough Rotation Shaft Rotation | 0
[7 empty] | 1 |  | 2 | [11 empty] | 3 | 0.4464% | 567 | Rough Jaw Actuation | 3
[7 empty] | 25 |  | 22 |  | 3 | [9 empty] | 50 | 7.44% | 567 | Stuck Blade | 50
[21 empty] | 0 | 0% | 567 | Blade does not pass front stop | 0
[8 empty] | 5 | 1 | 1 | 1 | [9 empty] | 2 | 0.2976% | 567 | 0.003 Jaw Gap Fail | 2
[8 empty] | 1 | 4 | 1 | [10 empty] | 4 | 0.5952% | 567 | 0.006 Jaw Gap Fail | 4
[8 empty] | 25 | 15 | 8 | 1 | [9 empty] | 16 | 2.381% | 567 | High Jaw Force | 16
[10 empty] | 1 | [10 empty] | 0 | 0% | 567 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 567 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 567 | Jaw Short | 0
[21 empty] | 0 | 0% | 567 | Fuse Switch | 0
[11 empty] | 1 | [9 empty] | 1 | 0.1488% | 567 | Continuity Fail | 1
[8 empty] | 22 | 1 | 6 | [10 empty] | 1 | 0.1488% | 567 | Failure To Lock/Unlock Trigger | 1
[21 empty] | 0 | 0% | 567 | Incorrect Blade Orientation | 0
[8 empty] | 22 |  | 2 | [10 empty] | 0 | 0% | 567 | Misassembled | 0
[8 empty] | 5 |  | 1 | [10 empty] | 0 | 0% | 567 | Missing Component | 0 | 2ea No Grease
[21 empty] | 0 | 0% | 567 | Cosmetic Defect | 0
[11 empty] | 1 | [8 empty] | 1 | 2 | 0.2976% | 567 | Damaged Harness | 2
[8 empty] | 2 | [11 empty] | 1 | 1 | 0.1488% | 567 | Scratched Logo | 1
[21 empty] | 0 | 0% | 567 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.2976% | 567 | Exposed Wire | 2
[21 empty] | 0 | 0% | 567 | Incorrect Hub Press | 0
[7 empty] | 9 | [13 empty] | 9 | 1.339% | 567 | Uneven Pin | 9
[7 empty] | 4 | [13 empty] | 4 | 0.5952% | 567 | Damaged actuation tube from nut setting | 4
[21 empty] | 0 | 0% | 567 | Incorrect Weld | 0
[7 empty] | 2 | [13 empty] | 2 | 0.2976% | 567 | Damaged Frame | 2
[7 empty] | 1 | [13 empty] | 1 | 0.1488% | 567 | Frames Weld Incorrectly | 1
[7 empty] | 3 | [13 empty] | 3 | 0.4464% | 567 | Misassembled Frames | 3 | 33ea Damaged Lower Jaw
[7 empty] | 1 | [13 empty] | 1 | 0.1488% | 567 | Misaligned Hub | 1 | 26ea Damaged Upper Jaw
[7 empty] | 1 | [13 empty] | 1 | 0.1488% | 567 | Jaws Do Not Open | 1
[21 empty] | 0 | 0% | 567 | Front Stop Not Touching | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1488% | 567 | Misassembled | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1488% | 567 | Defective Component | 1
[6 empty] | Total | 51 | 96 | 45 | 21 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 105 | 15.62% | 567
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1491104 | EB240 | 576 | 631 | 551 | 87.32% | 2023-06-05 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 18 | [12 empty] | 0 | 0% | 551 | Rough Rotation Shaft Rotation | 0
[7 empty] | 6 |  | 1 |  | 1 | [9 empty] | 8 | 1.268% | 551 | Rough Jaw Actuation | 8
[7 empty] | 16 |  | 6 |  | 3 | [9 empty] | 25 | 3.962% | 551 | Stuck Blade | 25
[21 empty] | 0 | 0% | 551 | Blade does not pass front stop | 0
[21 empty] | 0 | 0% | 551 | 0.003 Jaw Gap Fail | 0
[8 empty] | 4 | 2 | [10 empty] | 1 | 3 | 0.4754% | 551 | 0.006 Jaw Gap Fail | 3
[8 empty] | 25 | 14 | 33 | 11 | [9 empty] | 25 | 3.962% | 551 | High Jaw Force | 25
[9 empty] | 1 | [11 empty] | 1 | 0.1585% | 551 | Low Jaw Force | 1
[21 empty] | 0 | 0% | 551 | Front Stop Not Touching | 0
[8 empty] | 1 | 1 | [10 empty] | 1 | 2 | 0.317% | 551 | Jaw Short | 2
[21 empty] | 0 | 0% | 551 | Fuse Switch | 0
[21 empty] | 0 | 0% | 551 | Continuity Fail | 0
[8 empty] | 10 | [12 empty] | 0 | 0% | 551 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 551 | Incorrect Blade Orientation | 0
[8 empty] | 3 |  |  | 1 | [9 empty] | 1 | 0.1585% | 551 | Misassembled | 1
[8 empty] | 6 | [12 empty] | 0 | 0% | 551 | Missing Component | 0 | 2ea No Grease
[21 empty] | 0 | 0% | 551 | Missing Harness | 0
[11 empty] | 5 | [9 empty] | 5 | 0.7924% | 551 | Damaged Harness | 5
[8 empty] | 4 | [11 empty] | 2 | 2 | 0.317% | 551 | Scratched Logo | 2
[8 empty] | 1 | [12 empty] | 0 | 0% | 551 | Stuck Button | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.317% | 551 | Exposed Wire | 2
[21 empty] | 0 | 0% | 551 | Incorrect Hub Press | 0
[21 empty] | 0 | 0% | 551 | Uneven Pin | 0
[7 empty] | 2 | [13 empty] | 2 | 0.317% | 551 | Damaged actuation tube from nut setting | 2
[21 empty] | 0 | 0% | 551 | Incorrect Weld | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1585% | 551 | Damaged Frame | 1
[21 empty] | 0 | 0% | 551 | Frames Weld Incorrectly | 0
[21 empty] | 0 | 0% | 551 | Misassembled Frames | 0
[7 empty] | 2 | [13 empty] | 2 | 0.317% | 551 | Misaligned Hub | 2
[7 empty] | 1 | [13 empty] | 1 | 0.1585% | 551 | Jaws Do Not Open | 1
[21 empty] | 0 | 0% | 551 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 551 | Misassembled | 0
[21 empty] | 0 | 0% | 551 | Defective Component | 0
[6 empty] | Total | 30 | 72 | 25 | 33 | 21 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 80 | 12.68% | 551
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1493482 | EB240 | 576 | 647 | 569 | 87.94% | 2023-06-07 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 22 | [12 empty] | 0 | 0% | 569 | Rough Rotation Shaft Rotation | 0
[7 empty] | 7 |  | 1 | [11 empty] | 8 | 1.236% | 569 | Rough Jaw Actuation | 8
[7 empty] | 41 |  | 5 | [11 empty] | 46 | 7.11% | 569 | Stuck Blade | 46
[21 empty] | 0 | 0% | 569 | Blade does not pass front stop | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 569 | 0.003 Jaw Gap Fail | 0
[8 empty] | 1 | 1 | [11 empty] | 1 | 0.1546% | 569 | 0.006 Jaw Gap Fail | 1
[8 empty] | 25 | 1 | 19 | 5 | [9 empty] | 6 | 0.9274% | 569 | High Jaw Force | 6
[21 empty] | 0 | 0% | 569 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 569 | Front Stop Not Touching | 0
[8 empty] | 3 | 2 | [11 empty] | 2 | 0.3091% | 569 | Jaw Short | 2
[21 empty] | 0 | 0% | 569 | Fuse Switch | 0
[8 empty] | 8 | 1 | [11 empty] | 1 | 0.1546% | 569 | Continuity Fail | 1
[8 empty] | 13 | [12 empty] | 0 | 0% | 569 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 569 | Incorrect Blade Orientation | 0
[8 empty] | 17 | [12 empty] | 0 | 0% | 569 | Misassembled | 0
[8 empty] | 6 | [12 empty] | 0 | 0% | 569 | Missing Component | 0 | 2ea No Grease
[21 empty] | 0 | 0% | 569 | Missing Harness | 0
[21 empty] | 0 | 0% | 569 | Damaged Harness | 0
[8 empty] | 2 | [11 empty] | 1 | 1 | 0.1546% | 569 | Scratched Logo | 1
[8 empty] | 1 | [12 empty] | 0 | 0% | 569 | Particulate Matter | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.3091% | 569 | Exposed Wire | 2
[21 empty] | 0 | 0% | 569 | Incorrect Hub Press | 0
[7 empty] | 2 | [13 empty] | 2 | 0.3091% | 569 | Rough Jaw | 2
[7 empty] | 1 | [13 empty] | 1 | 0.1546% | 569 | Damaged actuation tube from nut setting | 1
[21 empty] | 0 | 0% | 569 | Incorrect Weld | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1546% | 569 | Damaged Frame | 1
[21 empty] | 0 | 0% | 569 | Frames Weld Incorrectly | 0
[7 empty] | 3 | [13 empty] | 3 | 0.4637% | 569 | Misassembled Frames | 3 | ea Damaged Lower Jaw
[7 empty] | 2 | [13 empty] | 2 | 0.3091% | 569 | Misaligned Hub | 2 | 5ea Damaged Upper Jaw
[7 empty] | 1 | [13 empty] | 1 | 0.1546% | 569 | Jaws Do Not Open | 1
[21 empty] | 0 | 0% | 569 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 569 | Misassembled | 0
[20 empty] | 1 | 1 | 0.1546% | 569 | Defective Component | 1
[6 empty] | Total | 60 | 99 | 11 | 19 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 78 | 12.06% | 569
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1492700 | EB240 | 576 | 660 | 569 | 86.21% | 2023-06-09 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 7 | 1 | 1 | [10 empty] | 1 | 0.1515% | 569 | Rough Rotation Shaft Rotation | 1
[7 empty] | 10 | [13 empty] | 10 | 1.515% | 569 | Rough Jaw Actuation | 10
[7 empty] | 50 |  | 5 | [11 empty] | 55 | 8.333% | 569 | Stuck Blade | 55
[21 empty] | 0 | 0% | 569 | Blade does not pass front stop | 0
[21 empty] | 0 | 0% | 569 | 0.003 Jaw Gap Fail | 0
[8 empty] | 2 | 2 | [10 empty] | 1 | 3 | 0.4545% | 569 | 0.006 Jaw Gap Fail | 3
[8 empty] | 9 |  | 6 | [10 empty] | 0 | 0% | 569 | High Jaw Force | 0
[21 empty] | 0 | 0% | 569 | Low Jaw Force | 0
[8 empty] | 2 | 2 | [11 empty] | 2 | 0.303% | 569 | Front Stop Not Touching | 2
[8 empty] | 1 | [12 empty] | 0 | 0% | 569 | Jaw Short | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 569 | Fuse Switch | 0
[8 empty] | 1 | [11 empty] | 1 | 1 | 0.1515% | 569 | Continuity Fail | 1
[8 empty] | 21 | [12 empty] | 0 | 0% | 569 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 569 | Incorrect Blade Orientation | 0
[8 empty] | 1 |  | 1 | [10 empty] | 0 | 0% | 569 | Misassembled | 0
[8 empty] | 4 | [11 empty] | 1 | 1 | 0.1515% | 569 | Missing Component | 1 | 3ea No Grease , QC: 1ea No Grease
[21 empty] | 0 | 0% | 569 | Missing Harness | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 569 | Particulate Matter | 0
[8 empty] | 1 | [12 empty] | 0 | 0% | 569 | Scratched Logo | 0
[8 empty] | 1 | [11 empty] | 4 | 4 | 0.6061% | 569 | Cosmetic Defect | 4
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 4 | [13 empty] | 4 | 0.6061% | 569 | Exposed Wire | 4
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 569 | Uneven Pins | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 569 | Uneven Jaws | 1
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 569 | Damaged actuation tube from nut setting | 1
[7 empty] | 2 | [13 empty] | 2 | 0.303% | 569 | Rough Jaw | 2
[21 empty] | 0 | 0% | 569 | Damaged Frame | 0
[21 empty] | 0 | 0% | 569 | Frames Weld Incorrectly | 0 | 2ea Damaged Tube
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 569 | Misassembled Frames | 1 | 2ea Damaged Lower Jaw
[21 empty] | 0 | 0% | 569 | Misaligned Hub | 0 | 5ea Damaged Upper Jaw
[7 empty] | 3 | [13 empty] | 3 | 0.4545% | 569 | Jaws Do Not Open | 3
[21 empty] | 0 | 0% | 569 | Front Stop Not Touching | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1515% | 569 | Misassembled | 1
[21 empty] | 0 | 0% | 569 | Defective Component | 0
[6 empty] | Total | 74 | 52 | 10 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 91 | 13.79% | 569
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1493483 | EB240 | 576 | 637 | 565 | 88.7% | 2023-06-15 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 5 |  |  | 2 | [9 empty] | 2 | 0.314% | 565 | Rough Rotation Shaft Rotation | 2
[7 empty] | 7 |  | 4 |  | 2 | [9 empty] | 13 | 2.041% | 565 | Rough Jaw Actuation | 13
[7 empty] | 29 |  | 6 |  | 1 | [9 empty] | 36 | 5.651% | 565 | Stuck Blade | 36
[21 empty] | 0 | 0% | 565 | Blade does not pass front stop | 0
[21 empty] | 0 | 0% | 565 | 0.003 Jaw Gap Fail | 0
[8 empty] | 4 | 3 | [11 empty] | 3 | 0.471% | 565 | 0.006 Jaw Gap Fail | 3
[8 empty] | 4 |  | 2 | [10 empty] | 0 | 0% | 565 | High Jaw Force | 0
[21 empty] | 0 | 0% | 565 | Low Jaw Force | 0
[20 empty] | 1 | 1 | 0.157% | 565 | Front Stop Not Touching | 1
[7 empty] | 1 |  |  |  | 1 | [9 empty] | 2 | 0.314% | 565 | Jaw Short | 2
[8 empty] | 1 |  |  | 1 | [9 empty] | 1 | 0.157% | 565 | Fuse Switch | 1
[21 empty] | 0 | 0% | 565 | Continuity Fail | 0
[8 empty] | 12 | [12 empty] | 0 | 0% | 565 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 565 | Incorrect Blade Orientation | 0
[8 empty] | 13 | [12 empty] | 0 | 0% | 565 | Misassembled | 0
[8 empty] | 7 | [12 empty] | 0 | 0% | 565 | Missing Component | 0 | 1ea No Grease
[11 empty] | 4 | [9 empty] | 4 | 0.6279% | 565 | Damaged Harness | 4
[21 empty] | 0 | 0% | 565 | Particulate Matter | 0
[21 empty] | 0 | 0% | 565 | Scratched Logo | 0
[8 empty] | 5 | [12 empty] | 0 | 0% | 565 | Cosmetic Defect | 0
[24 empty] | Shaft Level Scrap | 0
[21 empty] | 0 | 0% | 565 | Exposed Wire | 0
[7 empty] | 2 | [13 empty] | 2 | 0.314% | 565 | Uneven Pins | 2
[21 empty] | 0 | 0% | 565 | Uneven Jaws | 0
[7 empty] | 1 | [13 empty] | 1 | 0.157% | 565 | Damaged actuation tube from nut setting | 1
[7 empty] | 4 | [13 empty] | 4 | 0.6279% | 565 | Rough Jaw | 4
[7 empty] | 1 | [13 empty] | 1 | 0.157% | 565 | Damaged Frame | 1
[7 empty] | 1 | [13 empty] | 1 | 0.157% | 565 | Frames Weld Incorrectly | 1
[21 empty] | 0 | 0% | 565 | Misassembled Frames | 0
[21 empty] | 0 | 0% | 565 | Misaligned Hub | 0 | 3ea Damaged Upper Jaw
[21 empty] | 0 | 0% | 565 | Jaws Do Not Open | 0
[21 empty] | 0 | 0% | 565 | Front Stop Not Touching | 0
[7 empty] | 1 | [13 empty] | 1 | 0.157% | 565 | Misassembled | 1
[7 empty] | 1 | [13 empty] | 1 | 0.157% | 565 | Defective Component | 1
[6 empty] | Total | 48 | 51 | 13 | 2 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 73 | 11.46% | 565
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1492701 | EB240 | 576 | 671 | 569 | 84.8% | 2023-06-19 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 36 | [12 empty] | 0 | 0% | 569 | Rough Rotation Shaft Rotation | 0
[7 empty] | 7 |  | 1 | [11 empty] | 8 | 1.192% | 569 | Rough Jaw Actuation | 8
[7 empty] | 69 |  | 6 | [11 empty] | 75 | 11.18% | 569 | Stuck Blade | 75
[21 empty] | 0 | 0% | 569 | Blade does not pass front stop | 0
[21 empty] | 0 | 0% | 569 | 0.003 Jaw Gap Fail | 0
[10 empty] | 1 | [9 empty] | 1 | 1 | 0.149% | 569 | 0.006 Jaw Gap Fail | 1
[8 empty] | 3 | 1 | 15 | [10 empty] | 1 | 0.149% | 569 | High Jaw Force | 1
[21 empty] | 0 | 0% | 569 | Low Jaw Force | 0
[10 empty] | 1 | [10 empty] | 0 | 0% | 569 | Front Stop Not Touching | 0
[8 empty] | 2 | 2 | 1 | [10 empty] | 2 | 0.2981% | 569 | Jaw Short | 2
[21 empty] | 0 | 0% | 569 | Fuse Switch | 0
[21 empty] | 0 | 0% | 569 | Continuity Fail | 0
[21 empty] | 0 | 0% | 569 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 569 | Incorrect Blade Orientation | 0
[8 empty] | 8 | [12 empty] | 0 | 0% | 569 | Misassembled | 0
[8 empty] | 4 |  | 1 | [10 empty] | 0 | 0% | 569 | Missing Component | 0 | 2ea No Grease
[7 empty] | 1 |  |  |  | 6 | [9 empty] | 7 | 1.043% | 569 | Damaged Harness | 7
[21 empty] | 0 | 0% | 569 | Particulate Matter | 0
[10 empty] | 1 | [10 empty] | 0 | 0% | 569 | Scratched Logo | 0
[21 empty] | 0 | 0% | 569 | Cosmetic Defect | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.2981% | 569 | Exposed Wire | 2
[21 empty] | 0 | 0% | 569 | Uneven Pins | 0
[21 empty] | 0 | 0% | 569 | Uneven Jaws | 0
[7 empty] | 1 | [13 empty] | 1 | 0.149% | 569 | Damaged actuation tube from nut setting | 1
[7 empty] | 3 | [13 empty] | 3 | 0.4471% | 569 | Rough Jaw | 3
[21 empty] | 0 | 0% | 569 | Damaged Frame | 0
[7 empty] | 1 | [13 empty] | 1 | 0.149% | 569 | Frames Weld Incorrectly | 1
[21 empty] | 0 | 0% | 569 | Misassembled Frames | 0
[21 empty] | 0 | 0% | 569 | Misaligned Hub | 0
[21 empty] | 0 | 0% | 569 | Jaws Do Not Open | 0
[21 empty] | 0 | 0% | 569 | Front Stop Not Touching | 0
[7 empty] | 1 | [13 empty] | 1 | 0.149% | 569 | Misassembled | 1
[21 empty] | 0 | 0% | 569 | Defective Component | 0
[6 empty] | Total | 85 | 53 | 10 | 20 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 102 | 15.2% | 569
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1494748 | EB240 | 576 | 646 | 570 | 88.24% | 2023-06-22 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 17 | [12 empty] | 0 | 0% | 570 | Rough Rotation Shaft Rotation | 0
[7 empty] | 3 |  | 1 | [11 empty] | 4 | 0.6192% | 570 | Rough Jaw Actuation | 4
[7 empty] | 45 |  | 1 | [11 empty] | 46 | 7.121% | 570 | Stuck Blade | 46
[21 empty] | 0 | 0% | 570 | Blade does not pass front stop | 0
[21 empty] | 0 | 0% | 570 | 0.003 Jaw Gap Fail | 0
[8 empty] | 2 | 2 |  | 1 | [9 empty] | 3 | 0.4644% | 570 | 0.006 Jaw Gap Fail | 3
[8 empty] | 9 | 5 | 2 | [10 empty] | 5 | 0.774% | 570 | High Jaw Force | 5
[21 empty] | 0 | 0% | 570 | Low Jaw Force | 0
[21 empty] | 0 | 0% | 570 | Front Stop Not Touching | 0
[20 empty] | 1 | 1 | 0.1548% | 570 | Jaw Short | 1
[21 empty] | 0 | 0% | 570 | Fuse Switch | 0
[20 empty] | 1 | 1 | 0.1548% | 570 | Continuity Fail | 1
[8 empty] | 3 | [12 empty] | 0 | 0% | 570 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 570 | Incorrect Blade Orientation | 0
[8 empty] | 1 |  | 2 | [10 empty] | 0 | 0% | 570 | Misassembled | 0
[8 empty] | 6 |  | 2 | [9 empty] | 1 | 1 | 0.1548% | 570 | Missing Component | 1 | 4ea No Grease
[7 empty] | 1 | [13 empty] | 1 | 0.1548% | 570 | Damaged Jaw | 1
[21 empty] | 0 | 0% | 570 | Particulate Matter | 0
[8 empty] | 3 | [12 empty] | 0 | 0% | 570 | Scratched Logo | 0
[8 empty] | 1 | [11 empty] | 2 | 2 | 0.3096% | 570 | Cosmetic Defect | 2
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 2 | [13 empty] | 2 | 0.3096% | 570 | Exposed Wire | 2
[7 empty] | 1 | [13 empty] | 1 | 0.1548% | 570 | Incorrect/Missing Weld | 1
[21 empty] | 0 | 0% | 570 | Uneven Jaws | 0
[21 empty] | 0 | 0% | 570 | Damaged actuation tube from nut setting | 0
[21 empty] | 0 | 0% | 570 | Rough Jaw | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1548% | 570 | Damaged Frame | 1
[21 empty] | 0 | 0% | 570 | Frames Weld Incorrectly | 0
[7 empty] | 3 | [13 empty] | 3 | 0.4644% | 570 | Misassembled Frames | 3 | 2ea Damaged Upper Jaw
[7 empty] | 4 | [13 empty] | 4 | 0.6192% | 570 | Misaligned Hub | 4 | 5ea Damaged Lower Jaw
[7 empty] | 1 | [13 empty] | 1 | 0.1548% | 570 | Jaws Do Not Open | 1
[21 empty] | 0 | 0% | 570 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 570 | Misassembled | 0
[21 empty] | 0 | 0% | 570 | Defective Component | 0
[6 empty] | Total | 61 | 42 | 9 | 6 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 76 | 11.76% | 570
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Initial build Inspection (Scrapped Shafts) | Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework (Same Shafts) | 5th Rework (New Shafts) | 5th Rework (Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Codes | Description | Total | Comments
1493484 | EB240 | 576 | 648 | 557 | 85.96% | 2023-06-27 | [17 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[8 empty] | 15 |  | 2 | [10 empty] | 0 | 0% | 557 | Rough Rotation Shaft Rotation | 0
[7 empty] | 2 |  | 1 | [10 empty] | 1 | 4 | 0.6173% | 557 | Rough Jaw Actuation | 4
[7 empty] | 53 |  | 7 |  | 1 | [9 empty] | 61 | 9.414% | 557 | Stuck Blade | 61
[11 empty] | 1 | [9 empty] | 1 | 0.1543% | 557 | Blade does not pass front stop | 1
[21 empty] | 0 | 0% | 557 | 0.003 Jaw Gap Fail | 0
[8 empty] | 5 | 1 | [11 empty] | 1 | 0.1543% | 557 | 0.006 Jaw Gap Fail | 1
[8 empty] | 32 | 2 | 21 | 7 | [9 empty] | 9 | 1.389% | 557 | High Jaw Force | 9
[21 empty] | 0 | 0% | 557 | Low Jaw Force | 0
[8 empty] | 3 | 1 | 1 | [9 empty] | 2 | 3 | 0.463% | 557 | Front Stop Not Touching | 3
[8 empty] | 2 |  | 1 | 1 | [8 empty] | 1 | 2 | 0.3086% | 557 | Jaw Short | 2
[8 empty] | 1 |  |  | 1 | [9 empty] | 1 | 0.1543% | 557 | Fuse Switch | 1
[11 empty] | 1 | [9 empty] | 1 | 0.1543% | 557 | Continuity Fail | 1
[8 empty] | 7 |  | 1 | [10 empty] | 0 | 0% | 557 | Failure To Lock/Unlock Trigger | 0
[21 empty] | 0 | 0% | 557 | Incorrect Blade Orientation | 0
[8 empty] | 21 |  | 4 | [10 empty] | 0 | 0% | 557 | Misassembled | 0
[8 empty] | 5 |  | 1 | [10 empty] | 0 | 0% | 557 | Missing Component | 0 | 3ea No Grease
[11 empty] | 3 | [9 empty] | 3 | 0.463% | 557 | Damaged Harness | 3
[8 empty] | 1 |  | 1 | [10 empty] | 0 | 0% | 557 | Stuck Button | 0
[8 empty] | 2 | [12 empty] | 0 | 0% | 557 | Scratched Logo | 0
[21 empty] | 0 | 0% | 557 | Cosmetic Defect | 0
[24 empty] | Shaft Level Scrap | 0
[7 empty] | 3 | [13 empty] | 3 | 0.463% | 557 | Exposed Wire | 3
[21 empty] | 0 | 0% | 557 | Incorrect/Missing Weld | 0
[21 empty] | 0 | 0% | 557 | Uneven Jaws | 0
[21 empty] | 0 | 0% | 557 | Damaged actuation tube from nut setting | 0
[21 empty] | 0 | 0% | 557 | Rough Jaw | 0
[21 empty] | 0 | 0% | 557 | Damaged Frame | 0
[7 empty] | 1 | [13 empty] | 1 | 0.1543% | 557 | Frames Weld Incorrectly | 1 | 1ea Damaged Tube
[21 empty] | 0 | 0% | 557 | Misassembled Frames | 0 | 1ea Damaged Upper Jaw
[7 empty] | 1 | [13 empty] | 1 | 0.1543% | 557 | Misaligned Hub | 1
[21 empty] | 0 | 0% | 557 | Jaws Do Not Open | 0
[21 empty] | 0 | 0% | 557 | Front Stop Not Touching | 0
[21 empty] | 0 | 0% | 557 | Misassembled | 0
[21 empty] | 0 | 0% | 557 | Defective Component | 0
[6 empty] | Total | 60 | 94 | 12 | 32 | 15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 91 | 14.04% | 557

## EB015-EB215 Graphs
EB015/EB215 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 225
[5 empty] | 0.003 Jaw Gap Fail | 40
[5 empty] | Insulation Damage | 39
[5 empty] | Jaw Short | 29
[5 empty] | Rough Actuation of Jaw | 23
[5 empty] | Incorrect Insulation Orientation | 12
[5 empty] | Continuity Fail | 5
[5 empty] | Misassembled | 4
[5 empty] | High Jaw Force | 4
[5 empty] | Front Stops Not Touching | 2
[5 empty] | Fuse Switch | 1
[5 empty] | Low Jaw Force | 1
[5 empty] | Rough Knob Actuation | 1
[5 empty] | Damaged Component | 0
[5 empty] | Damaged Jaw | 0
[5 empty] | Middle Stops Not Touching | 0
Shop Order's Yield |  |  |  |  | Blade Activation | 0
Shop Order | Build QTY | Yield | Yield % | Date | Squeaky Blade | 0
1486334 | 2095 | 1871 | 89.31% | 2023-05-30 | Rough Trigger | 0
1488957 | 2098 | 1860 | 88.66% | 2023-06-01 | Incorrect knob collar press | 0
1492707 | 2094 | 1878 | 89.68% | 2023-06-06 | Both Not Touching | 0
1492705 | 2068 | 1872 | 90.52% | 2023-06-13 | Loose Knob Collar | 0
1492708 | 2032 | 1899 | 93.45% | 2023-06-23 | 0.005 Jaw Gap Fail | 0
1492706 | 2058 | 1880 | 91.35% | 2023-06-27 | Damaged Handle | 0
Average Yield: |  |  | 90.5% |  | Discolored Jaws | 0

## EB015-EB215
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486335 | EB215 | 1920 | 2133 | 1818 | 85.23% | 2023-03-31 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 117 |  | 2 | [8 empty] | 51 | 170 | 7.97% | 2133 | Stuck Blade | 170 | **Laser Welding Shim: .007 Shim**
[7 empty] | 5 | [11 empty] | 5 | 0.2344% | 2133 | Rough Actuation of Jaw | 5 | **Nut Setting Shim: .010 shim
[7 empty] | 58 | [10 empty] | 23 | 81 | 3.797% | 2133 | 0.003 Jaw Gap Fail | 81
[7 empty] | 2 |  | 2 | [9 empty] | 4 | 0.1875% | 2133 | 0.006 Jaw Gap Fail | 4
[7 empty] | 3 | [11 empty] | 3 | 0.1406% | 2133 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 2133 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2133 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2133 | Blade Does Not Pass Front Stop | 0
[7 empty] | 5 | [10 empty] | 3 | 8 | 0.3751% | 2133 | Jaw Short | 8
[7 empty] | 3 |  | 1 | [9 empty] | 4 | 0.1875% | 2133 | Insulation Damage | 4
[7 empty] | 2 | [11 empty] | 2 | 0.09376% | 2133 | Incorrect Insulation Orientation | 2
[9 empty] | 6 | [9 empty] | 6 | 0.2813% | 2133 | Damaged Harness | 6
[19 empty] | 0 | 0% | 2133 | Incorrect/Missing Weld | 0
[19 empty] | 0 | 0% | 2133 | Knob Collar Press incorrectly | 0
[17 empty] | 2 |  | 2 | 0.09376% | 2133 | Partial Pack | 2
[8 empty] | 8 | [10 empty] | 0 | 0% | 2133 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2133 | Particulate Matter | 0
[9 empty] | 1 | [8 empty] | 2 | 3 | 0.1406% | 2133 | Continuity Fail | 3
[8 empty] | 5 | 1 | [9 empty] | 1 | 0.04688% | 2133 | High Jaw Force | 1
[8 empty] | 4 | [10 empty] | 0 | 0% | 2133 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2133 | Return Spring Pop Out | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 2133 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2133 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 7ea
[9 empty] | 2 | [9 empty] | 2 | 0.09376% | 2133 | Failure To Unblock Trigger | 2 | Damaged Front Blade Pusher: 122ea
[8 empty] | 9 | 3 | [9 empty] | 3 | 0.1406% | 2133 | Misassembled | 3
[8 empty] | 1 | 1 | [9 empty] | 1 | 0.04688% | 2133 | Cosmetic Defect | 1
[19 empty] | 0 | 0% | 2133 | Missing Component | 0
[8 empty] | 1 | [9 empty] | 2 | 2 | 0.09376% | 2133 | Scratched Logo | 2
0
[19 empty] | 0 | 0% | 2133 | Insulation Damage | 0 | 9ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2133 | Incorrect Jaw Crimp | 0 | 3ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 1 | [11 empty] | 1 | 0.04688% | 2133 | Contamination | 1 | 19ea Contamination Pull Tube
[7 empty] | 4 | [11 empty] | 4 | 0.1875% | 2133 | Exposed Wire | 4 | 6ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2133 | Damaged Nut Thread | 0 | 5ea Contamination Lower Jaw
[19 empty] | 0 | 0% | 2133 | Incorrect/Missing Weld | 0
[19 empty] | 0 | 0% | 2133 | Damaged Blade | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04688% | 2133 | Stuck Blade | 1
[7 empty] | 12 | [11 empty] | 12 | 0.5626% | 2133 | Blade Pusher Coming Out | 12
[6 empty] | Total | 213 | 31 | 19 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 81 | 315 | 14.77% | 2133
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486327 | EB215 | 1920 | 2096 | 1827 | 87.17% | 2023-04-07 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 58 |  | 18 | [8 empty] | 31 | 107 | 5.105% | 2096 | Stuck Blade | 107 | **Laser Welding Shim: .007 Shim**
[7 empty] | 4 | [11 empty] | 4 | 0.1908% | 2096 | Rough Actuation of Jaw | 4 | **Nut Setting Shim: .010 shim
[7 empty] | 51 |  | 2 | [8 empty] | 21 | 74 | 3.531% | 2096 | 0.003 Jaw Gap Fail | 74
[7 empty] | 4 | [10 empty] | 1 | 5 | 0.2385% | 2096 | 0.006 Jaw Gap Fail | 5
[7 empty] | 3 | [11 empty] | 3 | 0.1431% | 2096 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 2096 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2096 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2096 | Blade Does Not Pass Front Stop | 0
[7 empty] | 4 | [11 empty] | 4 | 0.1908% | 2096 | Jaw Short | 4
[7 empty] | 11 | [10 empty] | 3 | 14 | 0.6679% | 2096 | Insulation Damage | 14
[7 empty] | 1 | [11 empty] | 1 | 0.04771% | 2096 | Incorrect Insulation Orientation | 1
[9 empty] | 4 | [9 empty] | 4 | 0.1908% | 2096 | Damaged Harness | 4
[19 empty] | 0 | 0% | 2096 | Partial Pack | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04771% | 2096 | Knob Collar Press incorrectly | 1
[9 empty] | 1 | [9 empty] | 1 | 0.04771% | 2096 | Contamination | 1
[8 empty] | 3 | 1 | [9 empty] | 1 | 0.04771% | 2096 | Rough Knob Actuation | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2096 | Particulate Matter | 0
[8 empty] | 2 | [9 empty] | 1 | 1 | 0.04771% | 2096 | Continuity Fail | 1
[8 empty] | 13 | 1 | [9 empty] | 1 | 0.04771% | 2096 | High Jaw Force | 1
[8 empty] | 4 | [10 empty] | 0 | 0% | 2096 | Low Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2096 | Return Spring Pop Out | 0
[8 empty] | 4 | 3 | [9 empty] | 3 | 0.1431% | 2096 | Fuse Switch | 3
[8 empty] | 1 | [10 empty] | 0 | 0% | 2096 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2096 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 135ea
[8 empty] | 16 | 3 | [9 empty] | 3 | 0.1431% | 2096 | Misassembled | 3
[8 empty] | 2 | [10 empty] | 0 | 0% | 2096 | Cosmetic Defect | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2096 | Missing Component | 0
[8 empty] | 5 | [9 empty] | 3 | 3 | 0.1431% | 2096 | Scratched Logo | 3
0
[19 empty] | 0 | 0% | 2096 | Insulation Damage | 0 | 4ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2096 | Incorrect Jaw Crimp | 0 | 8ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2096 | Contamination | 0 | 3ea Damaged Pull Tube
[7 empty] | 4 | [11 empty] | 4 | 0.1908% | 2096 | Exposed Wire | 4 | 2ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2096 | Damaged Nut Thread | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04771% | 2096 | Incorrect/Missing Weld | 1
[19 empty] | 0 | 0% | 2096 | Damaged Blade | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1431% | 2096 | Stuck Blade | 3
[7 empty] | 30 | [11 empty] | 30 | 1.431% | 2096 | Blade Pusher Coming Out | 30
[6 empty] | Total | 175 | 53 | 33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 60 | 268 | 12.79% | 2096
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486741 | EB215 | 1920 | 2012 | 1862 | 92.54% | 2023-04-08 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 12 |  | 2 | [8 empty] | 4 | 18 | 0.8946% | 2012 | Stuck Blade | 18 | **Laser Welding Shim: .007 Shim**
[7 empty] | 5 | [10 empty] | 2 | 7 | 0.3479% | 2012 | Rough Actuation of Jaw | 7 | **Nut Setting Shim: .010 shim
[7 empty] | 19 |  | 1 | [8 empty] | 11 | 31 | 1.541% | 2012 | 0.003 Jaw Gap Fail | 31
[9 empty] | 3 | [9 empty] | 3 | 0.1491% | 2012 | 0.006 Jaw Gap Fail | 3
[7 empty] | 1 |  | 1 | [9 empty] | 2 | 0.0994% | 2012 | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 2012 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2012 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 |  | 1 | [9 empty] | 2 | 0.0994% | 2012 | Blade Does Not Pass Front Stop | 2
[7 empty] | 7 | [10 empty] | 3 | 10 | 0.497% | 2012 | Jaw Short | 10
[7 empty] | 4 |  | 1 | [9 empty] | 5 | 0.2485% | 2012 | Insulation Damage | 5
[7 empty] | 3 | [11 empty] | 3 | 0.1491% | 2012 | Incorrect Insulation Orientation | 3
[9 empty] | 13 | [9 empty] | 13 | 0.6461% | 2012 | Damaged Harness | 13
[19 empty] | 0 | 0% | 2012 | Incorrect/Missing Weld | 0
[7 empty] | 2 |  | 1 | [9 empty] | 3 | 0.1491% | 2012 | Knob Collar Press incorrectly | 3
[19 empty] | 0 | 0% | 2012 | Partial Pack | 0
[19 empty] | 0 | 0% | 2012 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2012 | Particulate Matter | 0
[8 empty] | 3 | [9 empty] | 1 | 1 | 0.0497% | 2012 | Continuity Fail | 1
[8 empty] | 4 | [10 empty] | 0 | 0% | 2012 | High Jaw Force | 0
[8 empty] | 23 | 11 | [9 empty] | 11 | 0.5467% | 2012 | Low Jaw Force | 11
[8 empty] | 1 | [10 empty] | 0 | 0% | 2012 | Return Spring Pop Out | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2012 | Fuse Switch | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2012 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 6ea
[19 empty] | 0 | 0% | 2012 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 152ea
[8 empty] | 18 | [10 empty] | 0 | 0% | 2012 | Misassembled | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2012 | Cosmetic Defect | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2012 | Missing Component | 0
[18 empty] | 3 | 3 | 0.1491% | 2012 | Scratched Logo | 3
0
[19 empty] | 0 | 0% | 2012 | Insulation Damage | 0 | 13ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2012 | Incorrect Jaw Crimp | 0 | 10ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2012 | Contamination | 0 | 5ea Damaged Pull Tube
[7 empty] | 8 | [11 empty] | 8 | 0.3976% | 2012 | Exposed Wire | 8
[7 empty] | 1 | [11 empty] | 1 | 0.0497% | 2012 | Damaged Nut Thread | 1
[7 empty] | 1 | [11 empty] | 1 | 0.0497% | 2012 | Incorrect/Missing Weld | 1
[7 empty] | 1 | [11 empty] | 1 | 0.0497% | 2012 | Damaged Blade | 1
[7 empty] | 2 | [11 empty] | 2 | 0.0994% | 2012 | Stuck Blade | 2
[7 empty] | 25 | [11 empty] | 25 | 1.243% | 2012 | Blade Pusher Coming Out | 25
[6 empty] | Total | 92 | 55 | 34 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 24 | 150 | 7.455% | 2012
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486328 | EB215 | 1920 | 2038 | 1836 | 90.09% | 2023-04-12 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 43 |  | 3 | [8 empty] | 40 | 86 | 4.22% | 2038 | Stuck Blade | 86 | **Laser Welding Shim: .007 Shim**
[7 empty] | 13 | [10 empty] | 1 | 14 | 0.6869% | 2038 | Rough Actuation of Jaw | 14 | **Nut Setting Shim: .010 shim
[7 empty] | 20 |  | 4 | [8 empty] | 10 | 34 | 1.668% | 2038 | 0.003 Jaw Gap Fail | 34
[7 empty] | 8 | [10 empty] | 3 | 11 | 0.5397% | 2038 | 0.006 Jaw Gap Fail | 11
[7 empty] | 5 | [11 empty] | 5 | 0.2453% | 2038 | Front Stops Not Touching | 5
[19 empty] | 0 | 0% | 2038 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2038 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04907% | 2038 | Blade Does Not Pass Front Stop | 1
[7 empty] | 8 | [10 empty] | 3 | 11 | 0.5397% | 2038 | Jaw Short | 11
[7 empty] | 2 | [9 empty] | 3 | 1 | 6 | 0.2944% | 2038 | Insulation Damage | 6
[7 empty] | 11 | [11 empty] | 11 | 0.5397% | 2038 | Incorrect Insulation Orientation | 11
[9 empty] | 7 | [9 empty] | 7 | 0.3435% | 2038 | Damaged Harness | 7
[19 empty] | 0 | 0% | 2038 | Incorrect/Missing Weld | 0
[19 empty] | 0 | 0% | 2038 | Knob Collar Press incorrectly | 0
[17 empty] | 2 |  | 2 | 0.09814% | 2038 | Partial Pack | 2
[8 empty] | 6 | 2 | [9 empty] | 2 | 0.09814% | 2038 | Rough Knob Actuation | 2
[8 empty] | 1 | [10 empty] | 0 | 0% | 2038 | Particulate Matter | 0
[8 empty] | 4 | 1 | [9 empty] | 1 | 0.04907% | 2038 | Continuity Fail | 1
[8 empty] | 28 | [10 empty] | 0 | 0% | 2038 | High Jaw Force | 0
[19 empty] | 0 | 0% | 2038 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2038 | Return Spring Pop Out | 0
[8 empty] | 6 | [10 empty] | 0 | 0% | 2038 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2038 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2038 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 190ea
[8 empty] | 13 | 1 | [9 empty] | 1 | 0.04907% | 2038 | Misassembled | 1
[8 empty] | 2 | [8 empty] | 3 |  | 3 | 0.1472% | 2038 | Cosmetic Defect | 3
[8 empty] | 2 | [10 empty] | 0 | 0% | 2038 | Missing Component | 0
[19 empty] | 0 | 0% | 2038 | Scratched Logo | 0
0
[19 empty] | 0 | 0% | 2038 | Insulation Damage | 0 | 12ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2038 | Incorrect Jaw Crimp | 0 | 5ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2038 | Contamination | 0 | 6ea Damaged Pull Tube
[7 empty] | 2 | [11 empty] | 2 | 0.09814% | 2038 | Exposed Wire | 2 | 175ea Contamination Insulation
[19 empty] | 0 | 0% | 2038 | Damaged Nut Thread | 0 | 5ea Defective Upper Jaw
[19 empty] | 0 | 0% | 2038 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04907% | 2038 | Damaged Blade | 1
[7 empty] | 2 | [11 empty] | 2 | 0.09814% | 2038 | Cosmetic Defect (Cover Tube) | 2
[7 empty] | 2 | [11 empty] | 2 | 0.09814% | 2038 | Blade Pusher Coming Out | 2
[6 empty] | Total | 118 | 62 | 18 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 58 | 202 | 9.912% | 2038
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486742 | EB215 | 1920 | 2068 | 1854 | 89.65% | 2023-04-14 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 42 | [10 empty] | 16 | 58 | 2.805% | 2068 | Stuck Blade | 58 | **Laser Welding Shim: .007 Shim**
[7 empty] | 11 |  | 2 | [8 empty] | 1 | 14 | 0.677% | 2068 | Rough Actuation of Jaw | 14 | **Nut Setting Shim: .010 shim
[7 empty] | 45 |  | 8 | [8 empty] | 7 | 60 | 2.901% | 2068 | 0.003 Jaw Gap Fail | 60
[7 empty] | 11 |  | 2 | [8 empty] | 6 | 19 | 0.9188% | 2068 | 0.006 Jaw Gap Fail | 19
[7 empty] | 4 | [11 empty] | 4 | 0.1934% | 2068 | Front Stops Not Touching | 4
[19 empty] | 0 | 0% | 2068 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2068 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2068 | Blade Does Not Pass Front Stop | 0
[7 empty] | 4 |  | 1 | [8 empty] | 1 | 6 | 0.2901% | 2068 | Jaw Short | 6
[7 empty] | 3 |  | 2 | [7 empty] | 2 | 3 | 10 | 0.4836% | 2068 | Insulation Damage | 10
[7 empty] | 3 | [11 empty] | 3 | 0.1451% | 2068 | Incorrect Insulation Orientation | 3
[19 empty] | 0 | 0% | 2068 | Damaged Harness | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04836% | 2068 | Incorrect/Missing Weld | 1
[19 empty] | 0 | 0% | 2068 | Knob Collar Press incorrectly | 0
[17 empty] | 4 |  | 4 | 0.1934% | 2068 | Partial Pack | 4
[8 empty] | 7 | [10 empty] | 0 | 0% | 2068 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2068 | Particulate Matter | 0
[8 empty] | 4 | 1 | [8 empty] | 1 | 2 | 0.09671% | 2068 | Continuity Fail | 2
[8 empty] | 10 | [10 empty] | 0 | 0% | 2068 | High Jaw Force | 0
[8 empty] | 12 | 3 | [9 empty] | 3 | 0.1451% | 2068 | Low Jaw Force | 3
[8 empty] | 1 | [10 empty] | 0 | 0% | 2068 | Return Spring Pop Out | 0
[8 empty] | 2 | 2 | [9 empty] | 2 | 0.09671% | 2068 | Fuse Switch | 2
[19 empty] | 0 | 0% | 2068 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2068 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 117ea
[8 empty] | 8 | 3 | [9 empty] | 3 | 0.1451% | 2068 | Misassembled | 3
[8 empty] | 2 | [10 empty] | 0 | 0% | 2068 | Cosmetic Defect | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2068 | Missing Component | 0
[18 empty] | 1 | 1 | 0.04836% | 2068 | Scratched Logo | 1
0
[7 empty] | 3 | [11 empty] | 3 | 0.1451% | 2068 | Insulation Damage | 3 | 12ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2068 | Incorrect Jaw Crimp | 0 | 2ea Damaged Lower Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.09671% | 2068 | Stuck Blade | 2 | 1ea Damaged Pull Tube
[7 empty] | 3 | [11 empty] | 3 | 0.1451% | 2068 | Exposed Wire | 3 | 154ea Overtripped Upper Jaw
[19 empty] | 0 | 0% | 2068 | Damaged Nut Thread | 0
[19 empty] | 0 | 0% | 2068 | Incorrect/Missing Weld | 0
[7 empty] | 5 | [11 empty] | 5 | 0.2418% | 2068 | Damaged Blade | 5
[19 empty] | 0 | 0% | 2068 | Cosmetic Defect | 0
[7 empty] | 11 | [11 empty] | 11 | 0.5319% | 2068 | Blade Pusher Coming Out | 11
[6 empty] | Total | 148 | 47 | 24 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 36 | 214 | 10.35% | 2068
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486743 | EB215 | 1920 | 2012 | 1851 | 92% | 2023-04-19 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
|  |  |  | 69 |  |  | 37 |  | 1 | [8 empty] | 38 | 76 | 3.777% | 2012 | Stuck Blade | 76 | **Laser Welding Shim: .007 Shim**
[7 empty] | 4 | [10 empty] | 3 | 7 | 0.3479% | 2012 | Rough Actuation of Jaw | 7 | **Nut Setting Shim: .010 shim
[7 empty] | 3 |  | 3 | [8 empty] | 4 | 10 | 0.497% | 2012 | 0.003 Jaw Gap Fail | 10
[7 empty] | 3 | [10 empty] | 2 | 5 | 0.2485% | 2012 | 0.006 Jaw Gap Fail | 5
[7 empty] | 3 | [11 empty] | 3 | 0.1491% | 2012 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 2012 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2012 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2012 | Blade Does Not Pass Front Stop | 0
[7 empty] | 5 | [10 empty] | 4 | 9 | 0.4473% | 2012 | Jaw Short | 9
[7 empty] | 8 | [11 empty] | 8 | 0.3976% | 2012 | Insulation Damage | 8
[7 empty] | 1 | [11 empty] | 1 | 0.0497% | 2012 | Incorrect Insulation Orientation | 1
[9 empty] | 2 | [9 empty] | 2 | 0.0994% | 2012 | Damaged Harness | 2
[19 empty] | 0 | 0% | 2012 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.0497% | 2012 | Knob Collar Press incorrectly | 1
[19 empty] | 0 | 0% | 2012 | Partial Pack | 0
[8 empty] | 11 | [10 empty] | 0 | 0% | 2012 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2012 | Particulate Matter | 0
[8 empty] | 5 | 3 | [8 empty] | 6 | 9 | 0.4473% | 2012 | Continuity Fail | 9
[8 empty] | 25 | [10 empty] | 0 | 0% | 2012 | High Jaw Force | 0
[8 empty] | 8 | 1 | [9 empty] | 1 | 0.0497% | 2012 | Low Jaw Force | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2012 | Return Spring Pop Out | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2012 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2012 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2012 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 97ea
[8 empty] | 15 | 2 | [9 empty] | 2 | 0.0994% | 2012 | Misassembled | 2
[8 empty] | 1 | [10 empty] | 0 | 0% | 2012 | Cosmetic Defect | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2012 | Missing Component | 0
[19 empty] | 0 | 0% | 2012 | Scratched Logo | 0
[23 empty] | 0 | 3ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2012 | Insulation Damage | 0 | 4ea Incorrect Hub Press
[7 empty] | 4 | [11 empty] | 4 | 0.1988% | 2012 | Incorrect Jaw Crimp | 4 | 4ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2012 | Stuck Blade | 0 | 4ea Damaged Pull Tube
[7 empty] | 3 | [11 empty] | 3 | 0.1491% | 2012 | Exposed Wire | 3 | 5ea Damaged Upper Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.0994% | 2012 | Damaged Nut Thread | 2
[19 empty] | 0 | 0% | 2012 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.0497% | 2012 | Damaged Blade | 1
[19 empty] | 0 | 0% | 2012 | Cosmetic Defect | 0
[7 empty] | 17 | [11 empty] | 17 | 0.8449% | 2012 | Blade Pusher Coming Out | 17
[6 empty] | Total | 92 | 70 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 57 | 161 | 8.002% | 2012
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486329 | EB215 | 1920 | 2046 | 1850 | 90.42% | 2023-04-21 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 73 |  | 7 | [8 empty] | 28 | 108 | 5.279% | 2046 | Stuck Blade | 108 | **Laser Welding Shim: .007 Shim**
[7 empty] | 7 | [11 empty] | 7 | 0.3421% | 2046 | Rough Actuation of Jaw | 7 | **Nut Setting Shim: .010 shim
[7 empty] | 15 | [10 empty] | 10 | 25 | 1.222% | 2046 | 0.003 Jaw Gap Fail | 25
[7 empty] | 5 | [11 empty] | 5 | 0.2444% | 2046 | 0.006 Jaw Gap Fail | 5
[7 empty] | 2 | [11 empty] | 2 | 0.09775% | 2046 | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 2046 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2046 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2046 | Blade Does Not Pass Front Stop | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1466% | 2046 | Jaw Short | 3
[7 empty] | 8 | [10 empty] | 4 | 12 | 0.5865% | 2046 | Insulation Damage | 12
[7 empty] | 1 | [11 empty] | 1 | 0.04888% | 2046 | Incorrect Insulation Orientation | 1
[19 empty] | 0 | 0% | 2046 | Damaged Harness | 0
[19 empty] | 0 | 0% | 2046 | Incorrect/Missing Weld | 0
[9 empty] | 1 | [9 empty] | 1 | 0.04888% | 2046 | Knob Collar Press incorrectly | 1
[19 empty] | 0 | 0% | 2046 | Partial Pack | 0
[8 empty] | 9 | 5 | [9 empty] | 5 | 0.2444% | 2046 | Rough Knob Actuation | 5
[19 empty] | 0 | 0% | 2046 | Particulate Matter | 0
[18 empty] | 3 | 3 | 0.1466% | 2046 | Continuity Fail | 3
[8 empty] | 3 | 1 | [9 empty] | 1 | 0.04888% | 2046 | High Jaw Force | 1
[8 empty] | 1 | 4 | [9 empty] | 4 | 0.1955% | 2046 | Low Jaw Force | 4
[19 empty] | 0 | 0% | 2046 | Return Spring Pop Out | 0
[9 empty] | 1 | [9 empty] | 1 | 0.04888% | 2046 | Fuse Switch | 1
[9 empty] | 1 | [9 empty] | 1 | 0.04888% | 2046 | Blade Lever Stop Pop Out | 1 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2046 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 195ea
[8 empty] | 5 | 2 | [9 empty] | 2 | 0.09775% | 2046 | Misassembled | 2
[8 empty] | 3 | [10 empty] | 0 | 0% | 2046 | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 2046 | Missing Component | 0
[18 empty] | 3 | 3 | 0.1466% | 2046 | Scratched Logo | 3
[23 empty] | 0 | 3ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2046 | Insulation Damage | 0 | 3ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2046 | Incorrect Jaw Crimp | 0 | 3ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04888% | 2046 | Stuck Blade | 1 | 23ea Damaged Upper Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.09775% | 2046 | Exposed Wire | 2
[19 empty] | 0 | 0% | 2046 | Damaged Nut Thread | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1466% | 2046 | Incorrect/Missing Weld | 3
[7 empty] | 2 | [11 empty] | 2 | 0.09775% | 2046 | Damaged Blade | 2
[19 empty] | 0 | 0% | 2046 | Cosmetic Defect | 0
[7 empty] | 4 | [11 empty] | 4 | 0.1955% | 2046 | Blade Pusher Coming Out | 4
[6 empty] | Total | 126 | 21 | 22 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 48 | 196 | 9.58% | 2046
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486330 | EB215 | 1920 | 2109 | 1845 | 87.48% | 2023-04-24 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 56 |  | 8 | [8 empty] | 33 | 97 | 4.599% | 2109 | Stuck Blade | 97 | **Laser Welding Shim: .007 Shim**
[7 empty] | 42 | [10 empty] | 4 | 46 | 2.181% | 2109 | Rough Actuation of Jaw | 46 | **Nut Setting Shim: .010 shim
[7 empty] | 26 |  | 3 | [8 empty] | 2 | 31 | 1.47% | 2109 | 0.003 Jaw Gap Fail | 31
[7 empty] | 7 | [10 empty] | 2 | 9 | 0.4267% | 2109 | 0.006 Jaw Gap Fail | 9
[7 empty] | 4 | [11 empty] | 4 | 0.1897% | 2109 | Front Stops Not Touching | 4
[19 empty] | 0 | 0% | 2109 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2109 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04742% | 2109 | Blade Does Not Pass Front Stop | 1
[7 empty] | 3 |  | 1 | [8 empty] | 5 | 9 | 0.4267% | 2109 | Jaw Short | 9
[7 empty] | 4 |  | 4 | [8 empty] | 1 | 9 | 0.4267% | 2109 | Insulation Damage | 9
[7 empty] | 21 | [11 empty] | 21 | 0.9957% | 2109 | Incorrect Insulation Orientation | 21
[9 empty] | 3 | [9 empty] | 3 | 0.1422% | 2109 | Damaged Harness | 3
[19 empty] | 0 | 0% | 2109 | Incorrect/Missing Weld | 0
[7 empty] | 4 | [11 empty] | 4 | 0.1897% | 2109 | Knob Collar Press incorrectly | 4
[19 empty] | 0 | 0% | 2109 | Partial Pack | 0
[8 empty] | 14 | [10 empty] | 0 | 0% | 2109 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2109 | Particulate Matter | 0
[8 empty] | 3 | [9 empty] | 3 | 3 | 0.1422% | 2109 | Continuity Fail | 3
[8 empty] | 24 | [10 empty] | 0 | 0% | 2109 | High Jaw Force | 0
[8 empty] | 2 | 1 | [9 empty] | 1 | 0.04742% | 2109 | Low Jaw Force | 1
[8 empty] | 4 | [10 empty] | 0 | 0% | 2109 | Return Spring Pop Out | 0
[9 empty] | 1 | [9 empty] | 1 | 0.04742% | 2109 | Fuse Switch | 1
[19 empty] | 0 | 0% | 2109 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2109 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 172ea
[8 empty] | 12 | 1 | [9 empty] | 1 | 0.04742% | 2109 | Misassembled | 1
[8 empty] | 2 | [10 empty] | 0 | 0% | 2109 | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 2109 | Missing Component | 0
[18 empty] | 3 | 3 | 0.1422% | 2109 | Scratched Logo | 3
[23 empty] | 0 | 4ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2109 | Insulation Damage | 0 | 6ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2109 | Incorrect Jaw Crimp | 0 | 4ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04742% | 2109 | Stuck Blade | 1 | 22ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04742% | 2109 | Exposed Wire | 1 | 92ea Damaged Insulation
[19 empty] | 0 | 0% | 2109 | Damaged Nut Thread | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1422% | 2109 | Incorrect/Missing Weld | 3
[7 empty] | 1 | [11 empty] | 1 | 0.04742% | 2109 | Damaged Blade | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04742% | 2109 | Contamination | 1
[7 empty] | 14 | [11 empty] | 14 | 0.6638% | 2109 | Blade Pusher Coming Out | 14
[6 empty] | Total | 189 | 61 | 22 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 53 | 264 | 12.52% | 2109
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486754 | EB215 | 1920 | 2038 | 1852 | 90.87% | 2023-04-27 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 59 |  | 3 | [8 empty] | 16 | 78 | 3.827% | 2038 | Stuck Blade | 78 | **Laser Welding Shim: .007 Shim**
[7 empty] | 8 |  | 1 | [9 empty] | 9 | 0.4416% | 2038 | Rough Actuation of Jaw | 9 | **Nut Setting Shim: .010 shim
[7 empty] | 21 | [10 empty] | 9 | 30 | 1.472% | 2038 | 0.003 Jaw Gap Fail | 30
[9 empty] | 1 | [8 empty] | 8 | 9 | 0.4416% | 2038 | 0.006 Jaw Gap Fail | 9
[7 empty] | 2 | [11 empty] | 2 | 0.09814% | 2038 | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 2038 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2038 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04907% | 2038 | Blade Does Not Pass Front Stop | 1
[7 empty] | 4 |  | 1 | [8 empty] | 3 | 8 | 0.3925% | 2038 | Jaw Short | 8
[7 empty] | 10 |  | 1 | [8 empty] | 4 | 15 | 0.736% | 2038 | Insulation Damage | 15
[7 empty] | 6 | [11 empty] | 6 | 0.2944% | 2038 | Incorrect Insulation Orientation | 6
[9 empty] | 8 | [9 empty] | 8 | 0.3925% | 2038 | Damaged Harness | 8
[19 empty] | 0 | 0% | 2038 | Incorrect/Missing Weld | 0
[9 empty] | 1 | [9 empty] | 1 | 0.04907% | 2038 | Knob Collar Press incorrectly | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04907% | 2038 | Contamination | 1
[8 empty] | 13 | [10 empty] | 0 | 0% | 2038 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2038 | Particulate Matter | 0
[8 empty] | 1 | 1 | [8 empty] | 3 | 4 | 0.1963% | 2038 | Continuity Fail | 4
[8 empty] | 12 | 1 | [9 empty] | 1 | 0.04907% | 2038 | High Jaw Force | 1
[19 empty] | 0 | 0% | 2038 | Low Jaw Force | 0
[8 empty] | 3 | 2 | [9 empty] | 2 | 0.09814% | 2038 | Return Spring Pop Out | 2
[8 empty] | 2 | [10 empty] | 0 | 0% | 2038 | Fuse Switch | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2038 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 4ea
[19 empty] | 0 | 0% | 2038 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 35ea
[8 empty] | 4 | [10 empty] | 0 | 0% | 2038 | Misassembled | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2038 | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 2038 | Missing Component | 0
[18 empty] | 4 | 4 | 0.1963% | 2038 | Scratched Logo | 4
[23 empty] | 0 | 5ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2038 | Misassembled | 0 | 9ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.04907% | 2038 | Incorrect Jaw Crimp | 1 | 3ea Damaged Pull Tube
[19 empty] | 0 | 0% | 2038 | Stuck Blade | 0 | 8ea Damaged Upper Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.09814% | 2038 | Exposed Wire | 2 | 3ea Damaged Blade
[19 empty] | 0 | 0% | 2038 | Damaged Nut Thread | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04907% | 2038 | Incorrect/Missing Weld | 1
[19 empty] | 0 | 0% | 2038 | Damaged Blade | 0
[19 empty] | 0 | 0% | 2038 | Contamination | 0
[7 empty] | 2 | [11 empty] | 2 | 0.09814% | 2038 | Blade Pusher Coming Out | 2
[6 empty] | Total | 118 | 40 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 47 | 185 | 9.078% | 2038
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486331 | EB215 | 1920 | 2064 | 1865 | 90.36% | 2023-04-29 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 46 |  | 7 | [8 empty] | 22 | 75 | 3.634% | 2064 | Stuck Blade | 75 | **Laser Welding Shim: .007 Shim**
[7 empty] | 39 | [10 empty] | 6 | 45 | 2.18% | 2064 | Rough Actuation of Jaw | 45 | **Nut Setting Shim: .010 shim
[7 empty] | 14 | [10 empty] | 6 | 20 | 0.969% | 2064 | 0.003 Jaw Gap Fail | 20
[7 empty] | 7 | [11 empty] | 7 | 0.3391% | 2064 | 0.006 Jaw Gap Fail | 7
[19 empty] | 0 | 0% | 2064 | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 2064 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2064 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04845% | 2064 | Blade Does Not Pass Front Stop | 1
[7 empty] | 2 | [11 empty] | 2 | 0.0969% | 2064 | Jaw Short | 2
[7 empty] | 13 |  | 1 | [8 empty] | 6 | 20 | 0.969% | 2064 | Insulation Damage | 20
[7 empty] | 13 | [11 empty] | 13 | 0.6298% | 2064 | Incorrect Insulation Orientation | 13
[9 empty] | 3 | [9 empty] | 3 | 0.1453% | 2064 | Damaged Harness | 3
[19 empty] | 0 | 0% | 2064 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04845% | 2064 | Knob Collar Press incorrectly | 1
[19 empty] | 0 | 0% | 2064 | Contamination | 0
[8 empty] | 9 | [10 empty] | 0 | 0% | 2064 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2064 | Particulate Matter | 0
[8 empty] | 1 | [9 empty] | 1 | 1 | 0.04845% | 2064 | Continuity Fail | 1
[8 empty] | 8 | [10 empty] | 0 | 0% | 2064 | High Jaw Force | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2064 | Low Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2064 | Return Spring Pop Out | 0
[8 empty] | 3 | [9 empty] | 1 | 1 | 0.04845% | 2064 | Fuse Switch | 1
[19 empty] | 0 | 0% | 2064 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 3ea
[19 empty] | 0 | 0% | 2064 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 121ea
[8 empty] | 14 | [10 empty] | 0 | 0% | 2064 | Misassembled | 0
[8 empty] | 6 | [10 empty] | 0 | 0% | 2064 | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 2064 | Missing Component | 0
[18 empty] | 2 | 2 | 0.0969% | 2064 | Scratched Logo | 2
[23 empty] | 0 | 14ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2064 | Misassembled | 0 | 7ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.04845% | 2064 | Incorrect Jaw Crimp | 1 | 7ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04845% | 2064 | Stuck Blade | 1 | 59ea Over Tripped Upper Jaw
[19 empty] | 0 | 0% | 2064 | Exposed Wire | 0 | 1ea Contamination Lowr Jaw
[19 empty] | 0 | 0% | 2064 | Damaged Nut Thread | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1453% | 2064 | Incorrect/Missing Weld | 3
[7 empty] | 2 | [11 empty] | 2 | 0.0969% | 2064 | Damaged Blade | 2
[19 empty] | 0 | 0% | 2064 | Contamination | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04845% | 2064 | Blade Pusher Coming Out | 1
[6 empty] | Total | 144 | 44 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 44 | 199 | 9.641% | 2064
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486761 | EB215 | 1920 | 2086 | 1848 | 88.59% | 2023-05-06 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 82 | [10 empty] | 23 | 105 | 5.034% | 2086 | Stuck Blade | 105 | **Laser Welding Shim: .007 Shim**
[7 empty] | 24 | [10 empty] | 3 | 27 | 1.294% | 2086 | Rough Actuation of Jaw | 27 | **Nut Setting Shim: .010 shim
[7 empty] | 14 | [10 empty] | 4 | 18 | 0.8629% | 2086 | 0.003 Jaw Gap Fail | 18
[7 empty] | 11 | [10 empty] | 13 | 24 | 1.151% | 2086 | 0.006 Jaw Gap Fail | 24
[7 empty] | 6 | [11 empty] | 6 | 0.2876% | 2086 | Front Stops Not Touching | 6
[19 empty] | 0 | 0% | 2086 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2086 | Front and Middle Stops Not Touching | 0
[7 empty] | 2 | [11 empty] | 2 | 0.09588% | 2086 | Blade Does Not Pass Front Stop | 2
[7 empty] | 9 |  | 1 | [8 empty] | 5 | 15 | 0.7191% | 2086 | Jaw Short | 15
[7 empty] | 6 | [9 empty] | 1 | 9 | 16 | 0.767% | 2086 | Insulation Damage | 16
[7 empty] | 7 | [9 empty] | u |  | 7 | 0.3356% | 2086 | Incorrect Insulation Orientation | 7
[9 empty] | 4 | [9 empty] | 4 | 0.1918% | 2086 | Damaged Harness | 4
[19 empty] | 0 | 0% | 2086 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04794% | 2086 | Knob Collar Press incorrectly | 1
[17 empty] | 2 |  | 2 | 0.09588% | 2086 | Partial Pack | 2
[8 empty] | 1 | [10 empty] | 0 | 0% | 2086 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2086 | Particulate Matter | 0
[8 empty] | 3 | [9 empty] | 4 | 4 | 0.1918% | 2086 | Continuity Fail | 4
[8 empty] | 7 | [10 empty] | 0 | 0% | 2086 | High Jaw Force | 0
[19 empty] | 0 | 0% | 2086 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2086 | Return Spring Pop Out | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2086 | Fuse Switch | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2086 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2086 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 112ea
[8 empty] | 10 | [10 empty] | 0 | 0% | 2086 | Misassembled | 0
[8 empty] | 13 | [10 empty] | 0 | 0% | 2086 | Cosmetic Defect | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2086 | Missing Component | 0
[18 empty] | 3 | 3 | 0.1438% | 2086 | Scratched Logo | 3
[23 empty] | 0 | 12ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2086 | Misassembled | 0 | 5ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.04794% | 2086 | Incorrect Jaw Crimp | 1 | 8ea Damaged Pull Tube
[19 empty] | 0 | 0% | 2086 | Stuck Blade | 0 | 10ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04794% | 2086 | Exposed Wire | 1
[19 empty] | 0 | 0% | 2086 | Damaged Nut Thread | 0
[7 empty] | 2 | [11 empty] | 2 | 0.09588% | 2086 | Incorrect/Missing Weld | 2
[19 empty] | 0 | 0% | 2086 | Damaged Blade | 0
[19 empty] | 0 | 0% | 2086 | Contamination | 0
[19 empty] | 0 | 0% | 2086 | Blade Pusher Coming Out | 0
[6 empty] | Total | 166 | 39 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 64 | 238 | 11.41% | 2086
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486332 | EB215 | 1920 | 2137 | 1862 | 87.13% | 2023-05-12 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 93 |  | 5 | [8 empty] | 24 | 122 | 5.709% | 2137 | Stuck Blade | 122 | **Laser Welding Shim: .007 Shim**
[7 empty] | 22 | [10 empty] | 2 | 24 | 1.123% | 2137 | Rough Actuation of Jaw | 24 | **Nut Setting Shim: .010 shim
[7 empty] | 23 |  | 2 | [8 empty] | 1 | 26 | 1.217% | 2137 | 0.003 Jaw Gap Fail | 26
[7 empty] | 5 |  | 1 | [8 empty] | 2 | 8 | 0.3744% | 2137 | 0.006 Jaw Gap Fail | 8
[19 empty] | 0 | 0% | 2137 | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 2137 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2137 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04679% | 2137 | Blade Does Not Pass Front Stop | 1
[7 empty] | 2 | [10 empty] | 4 | 6 | 0.2808% | 2137 | Jaw Short | 6
[7 empty] | 14 |  | 3 | [9 empty] | 17 | 0.7955% | 2137 | Insulation Damage | 17
[7 empty] | 23 | [10 empty] | 3 | 26 | 1.217% | 2137 | Incorrect Insulation Orientation | 26
[19 empty] | 0 | 0% | 2137 | Damaged Harness | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04679% | 2137 | Incorrect/Missing Weld | 1
[7 empty] | 3 | [11 empty] | 3 | 0.1404% | 2137 | Knob Collar Press incorrectly | 3
[19 empty] | 0 | 0% | 2137 | Contamination | 0
[8 empty] | 5 | [10 empty] | 0 | 0% | 2137 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2137 | Particulate Matter | 0
[8 empty] | 3 | 2 | [8 empty] | 2 | 4 | 0.1872% | 2137 | Continuity Fail | 4
[8 empty] | 6 | 3 | [9 empty] | 3 | 0.1404% | 2137 | High Jaw Force | 3
[8 empty] | 3 | [10 empty] | 0 | 0% | 2137 | Low Jaw Force | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 2137 | Return Spring Pop Out | 0
[8 empty] | 1 | 1 | [9 empty] | 1 | 0.04679% | 2137 | Fuse Switch | 1
[19 empty] | 0 | 0% | 2137 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 19ea
[19 empty] | 0 | 0% | 2137 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 82ea
[8 empty] | 11 | 1 | [9 empty] | 1 | 0.04679% | 2137 | Misassembled | 1
[8 empty] | 9 | [10 empty] | 0 | 0% | 2137 | Cosmetic Defect | 0
[9 empty] | 2 | [9 empty] | 2 | 0.09359% | 2137 | Missing Component | 2
[19 empty] | 0 | 0% | 2137 | Scratched Logo | 0
[23 empty] | 0 | 19ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 2 | [11 empty] | 2 | 0.09359% | 2137 | Insulation Damage | 2 | 2ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2137 | Incorrect Jaw Crimp | 0 | 8ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04679% | 2137 | Stuck Blade | 1 | 25ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04679% | 2137 | Exposed Wire | 1 | 1ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2137 | Damaged Nut Thread | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04679% | 2137 | Incorrect/Missing Weld | 1
[19 empty] | 0 | 0% | 2137 | Damaged Blade | 0
[7 empty] | 5 | [11 empty] | 5 | 0.234% | 2137 | Missing Jaw Crimp | 5
[7 empty] | 20 | [11 empty] | 20 | 0.9359% | 2137 | Blade Pusher Coming Out | 20
[6 empty] | Total | 217 | 41 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 38 | 275 | 12.87% | 2137
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486333 | EB215 | 1920 | 2181 | 1879 | 86.15% | 2023-05-17 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 59 |  | 4 | [8 empty] | 11 | 74 | 3.393% | 2181 | Stuck Blade | 74 | **Laser Welding Shim: .007 Shim**
[7 empty] | 63 | [10 empty] | 3 | 66 | 3.026% | 2181 | Rough Actuation of Jaw | 66 | **Nut Setting Shim: .010 shim
[7 empty] | 24 | [10 empty] | 2 | 26 | 1.192% | 2181 | 0.003 Jaw Gap Fail | 26
[7 empty] | 1 | [11 empty] | 1 | 0.04585% | 2181 | 0.006 Jaw Gap Fail | 1
[7 empty] | 3 | [11 empty] | 3 | 0.1376% | 2181 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 2181 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2181 | Front and Middle Stops Not Touching | 0
[9 empty] | 1 | [9 empty] | 1 | 0.04585% | 2181 | Blade Does Not Pass Front Stop | 1
[7 empty] | 4 | [10 empty] | 1 | 5 | 0.2293% | 2181 | Jaw Short | 5
[7 empty] | 11 | [11 empty] | 11 | 0.5044% | 2181 | Insulation Damage | 11
[7 empty] | 36 |  | 3 | [8 empty] | 6 | 45 | 2.063% | 2181 | Incorrect Insulation Orientation | 45
[9 empty] | 2 | [9 empty] | 2 | 0.0917% | 2181 | Damaged Harness | 2
[19 empty] | 0 | 0% | 2181 | Incorrect/Missing Weld | 0
[7 empty] | 4 | [11 empty] | 4 | 0.1834% | 2181 | Knob Collar Press incorrectly | 4
[19 empty] | 0 | 0% | 2181 | Contamination | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2181 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2181 | Particulate Matter | 0
[8 empty] | 1 | 1 | [8 empty] | 1 | 2 | 0.0917% | 2181 | Continuity Fail | 2
[8 empty] | 16 | [10 empty] | 0 | 0% | 2181 | High Jaw Force | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2181 | Low Jaw Force | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 2181 | Return Spring Pop Out | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2181 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2181 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2181 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 144ea
[8 empty] | 13 | 6 | [9 empty] | 6 | 0.2751% | 2181 | Misassembled | 6
[8 empty] | 9 | [10 empty] | 0 | 0% | 2181 | Cosmetic Defect | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2181 | Missing Component | 0
[18 empty] | 1 | 1 | 0.04585% | 2181 | Scratched Logo | 1
[23 empty] | 0 | 4ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2181 | Insulation Damage | 0 | 3ea Damaged Pull Tube
[7 empty] | 3 | [11 empty] | 3 | 0.1376% | 2181 | Incorrect Jaw Crimp | 3 | 2ea Damaged Upper Jaw
[7 empty] | 3 | [11 empty] | 3 | 0.1376% | 2181 | Stuck Blade | 3
[7 empty] | 5 | [11 empty] | 5 | 0.2293% | 2181 | Exposed Wire | 5
[19 empty] | 0 | 0% | 2181 | Damaged Nut Thread | 0
[19 empty] | 0 | 0% | 2181 | Incorrect/Missing Weld | 0
[19 empty] | 0 | 0% | 2181 | Damaged Blade | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04585% | 2181 | Contamination | 1
[7 empty] | 44 | [11 empty] | 44 | 2.017% | 2181 | Blade Pusher Coming Out | 44
[6 empty] | Total | 261 | 55 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 25 | 303 | 13.89% | 2181
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486762 | EB215 | 1920 | 2157 | 1872 | 86.79% | 2023-05-18 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 57 | [10 empty] | 20 | 77 | 3.57% | 2157 | Stuck Blade | 77 | **Laser Welding Shim: .007 Shim**
[7 empty] | 50 | [10 empty] | 3 | 53 | 2.457% | 2157 | Rough Actuation of Jaw | 53 | **Nut Setting Shim: .010 shim
[7 empty] | 11 | [11 empty] | 11 | 0.51% | 2157 | 0.003 Jaw Gap Fail | 11
[7 empty] | 12 | [10 empty] | 2 | 14 | 0.649% | 2157 | 0.006 Jaw Gap Fail | 14
[7 empty] | 3 | [11 empty] | 3 | 0.1391% | 2157 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 2157 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2157 | Front and Middle Stops Not Touching | 0
[7 empty] | 2 | [11 empty] | 2 | 0.09272% | 2157 | Blade Does Not Pass Front Stop | 2
[7 empty] | 3 | [10 empty] | 4 | 7 | 0.3245% | 2157 | Jaw Short | 7
[7 empty] | 6 | [11 empty] | 6 | 0.2782% | 2157 | Insulation Damage | 6
[7 empty] | 27 | [10 empty] | 8 | 35 | 1.623% | 2157 | Incorrect Insulation Orientation | 35
[9 empty] | 3 | [9 empty] | 3 | 0.1391% | 2157 | Damaged Harness | 3
[17 empty] | 4 |  | 4 | 0.1854% | 2157 | Partial Pack | 4
[7 empty] | 1 | [11 empty] | 1 | 0.04636% | 2157 | Knob Collar Press incorrectly | 1
[7 empty] | 1 | [9 empty] | 1 |  | 2 | 0.09272% | 2157 | Scratched Logo | 2
[8 empty] | 3 | [10 empty] | 0 | 0% | 2157 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2157 | Particulate Matter | 0
[8 empty] | 2 | [9 empty] | 1 | 1 | 0.04636% | 2157 | Continuity Fail | 1
[8 empty] | 25 | [10 empty] | 0 | 0% | 2157 | High Jaw Force | 0
[19 empty] | 0 | 0% | 2157 | Low Jaw Force | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2157 | Return Spring Pop Out | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2157 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2157 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 2157 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 127ea
[8 empty] | 24 | [10 empty] | 0 | 0% | 2157 | Misassembled | 0
[8 empty] | 7 | [10 empty] | 0 | 0% | 2157 | Cosmetic Defect | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2157 | Missing Component | 0
[18 empty] | 2 | 2 | 0.09272% | 2157 | Scratched Logo | 2 | 5ea Double Blade (No Jaws)
[23 empty] | 0 | 11ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 1 | [11 empty] | 1 | 0.04636% | 2157 | Damaged Blade Pusher Rear | 1 | 12ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.04636% | 2157 | Incorrect Jaw Crimp | 1 | 16ea Damaged Pull Tube
[7 empty] | 6 | [11 empty] | 6 | 0.2782% | 2157 | Stuck Blade | 6 | 4ea Damaged Upper Jaw
[7 empty] | 9 | [11 empty] | 9 | 0.4172% | 2157 | Exposed Wire | 9
[7 empty] | 2 | [11 empty] | 2 | 0.09272% | 2157 | Incorrect Blade Orientation | 2
[7 empty] | 1 | [11 empty] | 1 | 0.04636% | 2157 | Incorrect/Missing Weld | 1
[7 empty] | 3 | [11 empty] | 3 | 0.1391% | 2157 | Damaged Blade | 3
[7 empty] | 1 | [11 empty] | 1 | 0.04636% | 2157 | Contamination | 1
[7 empty] | 40 | [11 empty] | 40 | 1.854% | 2157 | Blade Pusher Coming Out | 40
[6 empty] | Total | 237 | 69 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 40 | 285 | 13.21% | 2157
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1489595 | EB215 | 416 | 418 | 389 | 93.06% | 2023-05-26 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 1 | [10 empty] | 1 | 2 | 0.4785% | 418 | Stuck Blade | 2 | **Laser Welding Shim: .007 Shim**
[7 empty] | 4 | [11 empty] | 4 | 0.9569% | 418 | Rough Actuation of Jaw | 4 | **Nut Setting Shim: .010 shim
[7 empty] | 2 | [11 empty] | 2 | 0.4785% | 418 | 0.003 Jaw Gap Fail | 2
[7 empty] | 1 | [10 empty] | 2 | 3 | 0.7177% | 418 | 0.006 Jaw Gap Fail | 3
[7 empty] | 3 | [10 empty] | 1 | 4 | 0.9569% | 418 | Front Stops Not Touching | 4
[19 empty] | 0 | 0% | 418 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 418 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 418 | Blade Does Not Pass Front Stop | 0
[18 empty] | 1 | 1 | 0.2392% | 418 | Jaw Short | 1
[7 empty] | 5 | [11 empty] | 5 | 1.196% | 418 | Insulation Damage | 5
[7 empty] | 7 | [11 empty] | 7 | 1.675% | 418 | Incorrect Insulation Orientation | 7
[19 empty] | 0 | 0% | 418 | Damaged Harness | 0 | 10ea Harness Transfer to S/O
[19 empty] | 0 | 0% | 418 | Incorrect/Missing Weld | 0
[19 empty] | 0 | 0% | 418 | Knob Collar Press incorrectly | 0
[19 empty] | 0 | 0% | 418 | Scratched Logo | 0
[7 empty] | 1 | [11 empty] | 1 | 0.2392% | 418 | Rough Knob Actuation | 1
[19 empty] | 0 | 0% | 418 | Particulate Matter | 0
[19 empty] | 0 | 0% | 418 | Continuity Fail | 0
[7 empty] | 2 | [11 empty] | 2 | 0.4785% | 418 | High Jaw Force | 2
[19 empty] | 0 | 0% | 418 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 418 | Return Spring Pop Out | 0
[19 empty] | 0 | 0% | 418 | Fuse Switch | 0
[19 empty] | 0 | 0% | 418 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: ea
[19 empty] | 0 | 0% | 418 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: ea
[19 empty] | 0 | 0% | 418 | Misassembled | 0
[7 empty] | 1 | [11 empty] | 1 | 0.2392% | 418 | Cosmetic Defect | 1
[19 empty] | 0 | 0% | 418 | Missing Component | 0
[19 empty] | 0 | 0% | 418 | Scratched Logo | 0
0
[19 empty] | 0 | 0% | 418 | Damaged Blade Pusher Rear | 0
[19 empty] | 0 | 0% | 418 | Incorrect Jaw Crimp | 0 | 8ea Damaged Pull Tube
[19 empty] | 0 | 0% | 418 | Stuck Blade | 0
[19 empty] | 0 | 0% | 418 | Exposed Wire | 0
[19 empty] | 0 | 0% | 418 | Incorrect Blade Orientation | 0
[19 empty] | 0 | 0% | 418 | Incorrect/Missing Weld | 0
[19 empty] | 0 | 0% | 418 | Damaged Blade | 0
[19 empty] | 0 | 0% | 418 | Contamination | 0
[7 empty] | 2 | [11 empty] | 2 | 0.4785% | 418 | Blade Pusher Coming Out | 2
[6 empty] | Total | 29 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 34 | 8.134% | 418
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1486334 | EB215 | 1920 | 2095 | 1871 | 89.31% | 2023-05-30 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 38 |  | 2 | [8 empty] | 12 | 52 | 2.482% | 2095 | Stuck Blade | 52 | **Laser Welding Shim: .007 Shim**
[7 empty] | 16 |  | 1 | [8 empty] | 6 | 23 | 1.098% | 2095 | Rough Actuation of Jaw | 23 | **Nut Setting Shim: .010 shim
[7 empty] | 20 |  | 2 | [8 empty] | 2 | 24 | 1.146% | 2095 | 0.003 Jaw Gap Fail | 24
[7 empty] | 15 |  | 1 | [8 empty] | 2 | 18 | 0.8592% | 2095 | 0.006 Jaw Gap Fail | 18
[7 empty] | 5 | [11 empty] | 5 | 0.2387% | 2095 | Front Stops Not Touching | 5
[19 empty] | 0 | 0% | 2095 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2095 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2095 | Blade Does Not Pass Front Stop | 0
[19 empty] | 0 | 0% | 2095 | Jaw Short | 0
[7 empty] | 18 |  | 1 | [9 empty] | 19 | 0.9069% | 2095 | Insulation Damage | 19
[7 empty] | 35 | [10 empty] | 5 | 40 | 1.909% | 2095 | Incorrect Insulation Orientation | 40
[9 empty] | 7 | [9 empty] | 7 | 0.3341% | 2095 | Damaged Harness | 7
[19 empty] | 0 | 0% | 2095 | Incorrect/Missing Weld | 0
[19 empty] | 0 | 0% | 2095 | Knob Collar Press incorrectly | 0
[19 empty] | 0 | 0% | 2095 | Scratched Logo | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2095 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2095 | Particulate Matter | 0
[8 empty] | 3 | 2 | [8 empty] | 1 | 3 | 0.1432% | 2095 | Continuity Fail | 3
[8 empty] | 150 | 1 | [9 empty] | 1 | 0.04773% | 2095 | High Jaw Force | 1
[19 empty] | 0 | 0% | 2095 | Low Jaw Force | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2095 | Return Spring Pop Out | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2095 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2095 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 4 ea
[19 empty] | 0 | 0% | 2095 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 141 ea
[8 empty] | 11 | 4 | [9 empty] | 4 | 0.1909% | 2095 | Misassembled | 4
[8 empty] | 5 | 1 | [9 empty] | 1 | 0.04773% | 2095 | Cosmetic Defect | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2095 | Missing Component | 0
[19 empty] | 0 | 0% | 2095 | Scratched Logo | 0
0
[19 empty] | 0 | 0% | 2095 | Damaged Blade Pusher Rear | 0
[19 empty] | 0 | 0% | 2095 | Incorrect Jaw Crimp | 0 | 10ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 2 | [11 empty] | 2 | 0.09547% | 2095 | Stuck Blade | 2 | 3ea Damaged Pull Tube
[7 empty] | 3 | [11 empty] | 3 | 0.1432% | 2095 | Exposed Wire | 3 | 22ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2095 | Incorrect Blade Orientation | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04773% | 2095 | Incorrect/Missing Weld | 1
[7 empty] | 2 | [11 empty] | 2 | 0.09547% | 2095 | Damaged Blade | 2
[19 empty] | 0 | 0% | 2095 | Contamination | 0
[7 empty] | 20 | [11 empty] | 20 | 0.9547% | 2095 | Blade Pusher Coming Out | 20
[6 empty] | Total | 175 | 177 | 22 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 28 | 225 | 10.74% | 2095
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1488957 | EB215 | 1920 | 2098 | 1860 | 88.66% | 2023-06-01 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 67 |  | 6 | [8 empty] | 27 | 100 | 4.766% | 2098 | Stuck Blade | 100 | **Laser Welding Shim: .007 Shim**
[7 empty] | 5 | [11 empty] | 5 | 0.2383% | 2098 | Rough Actuation of Jaw | 5 | **Nut Setting Shim: .010 shim
[7 empty] | 13 |  | 1 | [9 empty] | 14 | 0.6673% | 2098 | 0.003 Jaw Gap Fail | 14
[7 empty] | 24 | [11 empty] | 24 | 1.144% | 2098 | 0.006 Jaw Gap Fail | 24
[7 empty] | 2 | [11 empty] | 2 | 0.09533% | 2098 | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 2098 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2098 | Front and Middle Stops Not Touching | 0
[7 empty] | 2 |  | 1 | [9 empty] | 3 | 0.143% | 2098 | Blade Does Not Pass Front Stop | 3
[7 empty] | 1 | [10 empty] | 1 | 2 | 0.09533% | 2098 | Jaw Short | 2
[7 empty] | 11 |  | 5 | [9 empty] | 16 | 0.7626% | 2098 | Insulation Damage | 16
[7 empty] | 26 |  | 5 | [8 empty] | 2 | 33 | 1.573% | 2098 | Incorrect Insulation Orientation | 33
[9 empty] | 4 | [9 empty] | 4 | 0.1907% | 2098 | Damaged Harness | 4
[19 empty] | 0 | 0% | 2098 | Incorrect/Missing Weld | 0
[7 empty] | 1 |  | 2 | [9 empty] | 3 | 0.143% | 2098 | Knob Collar Press incorrectly | 3
[19 empty] | 0 | 0% | 2098 | Scratched Logo | 0
[8 empty] | 5 | [10 empty] | 0 | 0% | 2098 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2098 | Particulate Matter | 0
[8 empty] | 2 | [9 empty] | 2 | 2 | 0.09533% | 2098 | Continuity Fail | 2
[8 empty] | 90 | 1 | [9 empty] | 1 | 0.04766% | 2098 | High Jaw Force | 1
[19 empty] | 0 | 0% | 2098 | Low Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2098 | Return Spring Pop Out | 0
[19 empty] | 0 | 0% | 2098 | Fuse Switch | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2098 | Blade Lever Stop Pop Out | 0 | BP Cap Separation: 11 ea
[19 empty] | 0 | 0% | 2098 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 38 ea
[8 empty] | 9 | [10 empty] | 0 | 0% | 2098 | Misassembled | 0
[8 empty] | 6 | 2 | [9 empty] | 2 | 0.09533% | 2098 | Cosmetic Defect | 2
[19 empty] | 0 | 0% | 2098 | Missing Component | 0
[19 empty] | 0 | 0% | 2098 | Scratched Logo | 0
0
[7 empty] | 2 | [11 empty] | 2 | 0.09533% | 2098 | Incorrect Blade Stop Crimp | 2
[19 empty] | 0 | 0% | 2098 | Incorrect Jaw Crimp | 0 | 4ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 1 | [11 empty] | 1 | 0.04766% | 2098 | Stuck Blade | 1 | 46ea Damaged Pull Tube (can not use per Eng.)
[7 empty] | 2 | [11 empty] | 2 | 0.09533% | 2098 | Exposed Wire | 2 | 6ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2098 | Incorrect Blade Orientation | 0 | 5ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04766% | 2098 | Incorrect/Missing Weld | 1
[7 empty] | 5 | [11 empty] | 5 | 0.2383% | 2098 | Damaged Blade | 5
[19 empty] | 0 | 0% | 2098 | Contamination | 0
[7 empty] | 15 | [11 empty] | 15 | 0.715% | 2098 | Blade Pusher Coming Out | 15
[6 empty] | Total | 178 | 114 | 27 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 32 | 237 | 11.3% | 2098
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1492707 | EB215 | 1920 | 2094 | 1878 | 89.68% | 2023-06-06 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 39 |  | 4 | [8 empty] | 17 | 60 | 2.865% | 2094 | Stuck Blade | 60 | **Laser Welding Shim: .007 Shim**
[7 empty] | 10 |  | 1 | [8 empty] | 1 | 12 | 0.5731% | 2094 | Rough Actuation of Jaw | 12 | **Nut Setting Shim: .010 shim
[7 empty] | 11 |  | 3 | [8 empty] | 2 | 16 | 0.7641% | 2094 | 0.003 Jaw Gap Fail | 16
[7 empty] | 5 |  | 1 | [9 empty] | 6 | 0.2865% | 2094 | 0.006 Jaw Gap Fail | 6
[7 empty] | 3 | [11 empty] | 3 | 0.1433% | 2094 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 2094 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2094 | Front and Middle Stops Not Touching | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1433% | 2094 | Blade Does Not Pass Front Stop | 3
[7 empty] | 3 |  | 1 | [8 empty] | 2 | 6 | 0.2865% | 2094 | Jaw Short | 6
[7 empty] | 11 | [9 empty] | 1 | 4 | 16 | 0.7641% | 2094 | Insulation Damage | 16
[7 empty] | 6 | [11 empty] | 6 | 0.2865% | 2094 | Incorrect Insulation Orientation | 6
[9 empty] | 3 | [9 empty] | 3 | 0.1433% | 2094 | Damaged Harness | 3
[19 empty] | 0 | 0% | 2094 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04776% | 2094 | Knob Collar Press incorrectly | 1
[19 empty] | 0 | 0% | 2094 | Scratched Logo | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2094 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2094 | Particulate Matter | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 2094 | Continuity Fail | 0
[8 empty] | 23 | [10 empty] | 0 | 0% | 2094 | High Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2094 | Low Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2094 | Return Spring Pop Out | 0
[19 empty] | 0 | 0% | 2094 | Fuse Switch | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2094 | Blade Lever Stop Pop Out | 0 | BP Cap Separation:
[19 empty] | 0 | 0% | 2094 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 88ea
[8 empty] | 9 | 1 | [9 empty] | 1 | 0.04776% | 2094 | Misassembled | 1
[8 empty] | 11 | 1 | [9 empty] | 1 | 0.04776% | 2094 | Cosmetic Defect | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2094 | Missing Component | 0
[19 empty] | 0 | 0% | 2094 | Scratched Logo | 0
0
[7 empty] | 2 | [11 empty] | 2 | 0.09551% | 2094 | Incorrect Blade Stop Crimp | 2
[19 empty] | 0 | 0% | 2094 | Incorrect Jaw Crimp | 0 | 2ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 2 | [11 empty] | 2 | 0.09551% | 2094 | Stuck Blade | 2 | 12ea Damaged Upper Jaw
[7 empty] | 3 | [11 empty] | 3 | 0.1433% | 2094 | Exposed Wire | 3
[7 empty] | 1 | [11 empty] | 1 | 0.04776% | 2094 | Defective componet | 1
[7 empty] | 2 | [11 empty] | 2 | 0.09551% | 2094 | Incorrect/Missing Weld | 2
[7 empty] | 10 | [11 empty] | 10 | 0.4776% | 2094 | Damaged Blade | 10
[7 empty] | 21 | [11 empty] | 21 | 1.003% | 2094 | Insulation Damage | 21
[7 empty] | 41 | [11 empty] | 41 | 1.958% | 2094 | Blade Pusher Coming Out | 41
[6 empty] | Total | 174 | 51 | 15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 26 | 216 | 10.32% | 2094
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1492705 | EB215 | 1920 | 2068 | 1872 | 90.52% | 2023-06-13 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 78 | [10 empty] | 28 | 106 | 5.126% | 2068 | Stuck Blade | 106 | **Laser Welding Shim: .007 Shim**
[7 empty] | 5 | [11 empty] | 5 | 0.2418% | 2068 | Rough Actuation of Jaw | 5 | **Nut Setting Shim: .010 shim
[7 empty] | 8 |  | 1 | [8 empty] | 1 | 10 | 0.4836% | 2068 | 0.003 Jaw Gap Fail | 10
[7 empty] | 1 | [11 empty] | 1 | 0.04836% | 2068 | 0.006 Jaw Gap Fail | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04836% | 2068 | Front Stops Not Touching | 1
[19 empty] | 0 | 0% | 2068 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2068 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04836% | 2068 | Blade Does Not Pass Front Stop | 1
[18 empty] | 1 | 1 | 0.04836% | 2068 | Jaw Short | 1
[7 empty] | 3 | [11 empty] | 3 | 0.1451% | 2068 | Insulation Damage | 3
[7 empty] | 2 | [10 empty] | 2 | 4 | 0.1934% | 2068 | Incorrect Insulation Orientation | 4
[9 empty] | 5 | [9 empty] | 5 | 0.2418% | 2068 | Damaged Harness | 5
[19 empty] | 0 | 0% | 2068 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04836% | 2068 | Knob Collar Press incorrectly | 1
[19 empty] | 0 | 0% | 2068 | Scratched Logo | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 2068 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2068 | Particulate Matter | 0
[8 empty] | 2 | 2 | [8 empty] | 1 | 3 | 0.1451% | 2068 | Continuity Fail | 3
[8 empty] | 9 | 2 | [9 empty] | 2 | 0.09671% | 2068 | High Jaw Force | 2
[9 empty] | 1 | [9 empty] | 1 | 0.04836% | 2068 | Low Jaw Force | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2068 | Return Spring Pop Out | 0
[18 empty] | 1 | 1 | 0.04836% | 2068 | Fuse Switch | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 2068 | Blade Lever Stop Pop Out | 0 | BP Cap Separation:
[19 empty] | 0 | 0% | 2068 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 95ea
[8 empty] | 8 | 2 | [9 empty] | 2 | 0.09671% | 2068 | Misassembled | 2 | Insulation Damage: 135
[8 empty] | 11 | [10 empty] | 0 | 0% | 2068 | Cosmetic Defect | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2068 | Missing Component | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2068 | Damaged Harness | 0
0
[7 empty] | 2 | [11 empty] | 2 | 0.09671% | 2068 | Incorrect Blade Stop Crimp | 2
[19 empty] | 0 | 0% | 2068 | Incorrect Jaw Crimp | 0 | 3ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 3 | [11 empty] | 3 | 0.1451% | 2068 | Stuck Blade | 3 | 51ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04836% | 2068 | Exposed Wire | 1 | 1ea Contamination (Lower Jaw)
[7 empty] | 2 | [11 empty] | 2 | 0.09671% | 2068 | Defective componet | 2 | 3ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2068 | Incorrect/Missing Weld | 0 | 5ea Damaged Pull Tube
[7 empty] | 6 | [11 empty] | 6 | 0.2901% | 2068 | Damaged Blade | 6
[7 empty] | 1 | [11 empty] | 1 | 0.04836% | 2068 | Insulation Damage | 1
[7 empty] | 33 | [11 empty] | 33 | 1.596% | 2068 | Blade Pusher Coming Out | 33
[6 empty] | Total | 148 | 37 | 13 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 34 | 195 | 9.429% | 2068
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1492708 | EB215 | 1920 | 2032 | 1899 | 93.45% | 2023-06-23 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 20 | [10 empty] | 3 | 23 | 1.132% | 2032 | Stuck Blade | 23 | **Laser Welding Shim: .007 Shim**
[7 empty] | 11 | [10 empty] | 1 | 12 | 0.5906% | 2032 | Rough Actuation of Jaw | 12 | **Nut Setting Shim: .010 shim
[7 empty] | 13 | [10 empty] | 1 | 14 | 0.689% | 2032 | 0.003 Jaw Gap Fail | 14
[7 empty] | 1 | [11 empty] | 1 | 0.04921% | 2032 | 0.006 Jaw Gap Fail | 1
[19 empty] | 0 | 0% | 2032 | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 2032 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2032 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2032 | Blade Does Not Pass Front Stop | 0
[7 empty] | 12 | [10 empty] | 5 | 17 | 0.8366% | 2032 | Jaw Short | 17
[7 empty] | 16 |  | 2 | [8 empty] | 1 | 19 | 0.935% | 2032 | Insulation Damage | 19
[7 empty] | 3 | [11 empty] | 3 | 0.1476% | 2032 | Incorrect Insulation Orientation | 3
[9 empty] | 5 | [9 empty] | 5 | 0.2461% | 2032 | Damaged Harness | 5
[19 empty] | 0 | 0% | 2032 | Missing Harness | 0
[7 empty] | 4 | [11 empty] | 4 | 0.1969% | 2032 | Knob Collar Press incorrectly | 4
[19 empty] | 0 | 0% | 2032 | Scratched Logo | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 2032 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2032 | Particulate Matter | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2032 | Continuity Fail | 0
[8 empty] | 38 | 1 | [9 empty] | 1 | 0.04921% | 2032 | High Jaw Force | 1
[8 empty] | 11 | [10 empty] | 0 | 0% | 2032 | Low Jaw Force | 0
[8 empty] | 3 | 1 | [9 empty] | 1 | 0.04921% | 2032 | Return Spring Pop Out | 1
[8 empty] | 3 | [10 empty] | 0 | 0% | 2032 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2032 | Blade Lever Stop Pop Out | 0 | BP Cap Separation:
[19 empty] | 0 | 0% | 2032 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 76ea
[8 empty] | 15 | [10 empty] | 0 | 0% | 2032 | Misassembled | 0
[8 empty] | 5 | [9 empty] | 1 | 1 | 0.04921% | 2032 | Cosmetic Defect | 1
[19 empty] | 0 | 0% | 2032 | Missing Component | 0
[19 empty] | 0 | 0% | 2032 | Damaged Harness | 0
0
[19 empty] | 0 | 0% | 2032 | Incorrect Blade Stop Crimp | 0 | 4ea Missing Revert Dimple Press (No JawS)
[7 empty] | 1 | [11 empty] | 1 | 0.04921% | 2032 | Incorrect Jaw Crimp | 1 | 4ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 2 | [11 empty] | 2 | 0.09843% | 2032 | Stuck Blade | 2 | 3ea Damaged Upper Jaw
[7 empty] | 1 | [11 empty] | 1 | 0.04921% | 2032 | Exposed Wire | 1
[19 empty] | 0 | 0% | 2032 | Defective componet | 0
[19 empty] | 0 | 0% | 2032 | Incorrect/Missing Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04921% | 2032 | Damaged Blade | 1
[7 empty] | 2 | [11 empty] | 2 | 0.09843% | 2032 | Defective Component | 2
[7 empty] | 25 | [11 empty] | 25 | 1.23% | 2032 | Blade Pusher Coming Out | 25
[6 empty] | Total | 112 | 80 | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 133 | 6.545% | 2032
| Shop Order | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description | Build QTY | Comments
1492706 | EB215 | 1920 | 2058 | 1880 | 91.35% | 2023-06-27 | [15 empty] | Top Level Non-Conformance | 578.5 | **Production Shop Order**
[7 empty] | 73 | [10 empty] | 18 | 91 | 4.422% | 2058 | Stuck Blade | 91 | **Laser Welding Shim: .007 Shim**
[7 empty] | 6 | [11 empty] | 6 | 0.2915% | 2058 | Rough Actuation of Jaw | 6 | **Nut Setting Shim: .010 shim
[7 empty] | 14 |  | 1 | [8 empty] | 1 | 16 | 0.7775% | 2058 | 0.003 Jaw Gap Fail | 16
[7 empty] | 1 | [10 empty] | 1 | 2 | 0.09718% | 2058 | 0.006 Jaw Gap Fail | 2
[7 empty] | 1 | [11 empty] | 1 | 0.04859% | 2058 | Front Stops Not Touching | 1
[19 empty] | 0 | 0% | 2058 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2058 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04859% | 2058 | Blade Does Not Pass Front Stop | 1
[7 empty] | 6 |  | 1 | [8 empty] | 4 | 11 | 0.5345% | 2058 | Jaw Short | 11
[7 empty] | 14 | [10 empty] | 2 | 16 | 0.7775% | 2058 | Insulation Damage | 16
[7 empty] | 4 | [10 empty] | 1 | 5 | 0.243% | 2058 | Incorrect Insulation Orientation | 5
[9 empty] | 9 | [9 empty] | 9 | 0.4373% | 2058 | Damaged Harness | 9
[19 empty] | 0 | 0% | 2058 | Missing Harness | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04859% | 2058 | Knob Collar Press incorrectly | 1
[19 empty] | 0 | 0% | 2058 | Scratched Logo | 0
[8 empty] | 2 | 1 | [9 empty] | 1 | 0.04859% | 2058 | Rough Knob Actuation | 1
[8 empty] | 2 | [10 empty] | 0 | 0% | 2058 | Particulate Matter | 0
[8 empty] | 3 | 2 | [9 empty] | 2 | 0.09718% | 2058 | Continuity Fail | 2
[8 empty] | 29 | 1 | [9 empty] | 1 | 0.04859% | 2058 | High Jaw Force | 1
[8 empty] | 2 | [10 empty] | 0 | 0% | 2058 | Low Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2058 | Return Spring Pop Out | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2058 | Fuse Switch | 0
[19 empty] | 0 | 0% | 2058 | Blade Lever Stop Pop Out | 0 | BP Cap Separation:
[19 empty] | 0 | 0% | 2058 | Failure To Unblock Trigger | 0 | Damaged Front Blade Pusher: 67ea
[8 empty] | 2 | [10 empty] | 0 | 0% | 2058 | Misassembled | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 2058 | Cosmetic Defect | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2058 | Missing Component | 0
[19 empty] | 0 | 0% | 2058 | Damaged Harness | 0
0
[7 empty] | 2 | [11 empty] | 2 | 0.09718% | 2058 | Misassembled | 2
[7 empty] | 1 | [11 empty] | 1 | 0.04859% | 2058 | Incorrect Jaw Crimp | 1 | 8ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 2058 | Stuck Blade | 0 | 28ea Damaged Upper Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.09718% | 2058 | Exposed Wire | 2 | 1ea Damaged Pull Tube
[19 empty] | 0 | 0% | 2058 | Defective componet | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04859% | 2058 | Incorrect/Missing Weld | 1
[19 empty] | 0 | 0% | 2058 | Damaged Blade | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04859% | 2058 | Contamination | 1
[7 empty] | 10 | [11 empty] | 10 | 0.4859% | 2058 | Blade Pusher Coming Out | 10
[6 empty] | Total | 138 | 46 | 15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 180 | 8.746% | 2058

## EB016-EB216 Graphs
EB016/EB216 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 126
[5 empty] | Rough Actuation of Jaw | 54
[5 empty] | Incorrect Insulation Orientation | 47
[5 empty] | 0.003 Jaw Gap Fail | 45
[5 empty] | Jaw Short | 19
[5 empty] | Insulation Damage | 13
[5 empty] | 0.006 Jaw Gap Fail | 12
[5 empty] | Continuity Fail | 5
[5 empty] | Misassembled | 4
[5 empty] | High Jaw Force | 3
[5 empty] | Front Stops Not Touching | 2
[5 empty] | Low Jaw Force | 1
[5 empty] | Fuse Switch | 0
[5 empty] | Rough Knob Actuation | 0
[5 empty] | Blade Activation | 0
[5 empty] | Damaged Component | 0
Shop Order's Yield |  |  |  |  | Rough Trigger | 0
Shop Order | Build QTY | Yield | Yield % | Date | Middle Stops Not Touching | 0
1486757 | 1192 | 1134 | 95.13% | 2023-04-18 | Squeaky Blade | 0
1486758 | 2067 | 1855 | 89.74% | 2023-05-02 | Damaged Jaw | 0
1486337 | 1249 | 1128 | 90.31% | 2023-05-09 | Both Stops Not Touching | 0
1488010 | 1259 | 1104 | 87.69% | 2023-05-19 | Damaged Handle | 0
1486338 | 1225 | 1122 | 91.59% | 2023-06-07 | Collapsed Front Stop | 0
| #N/A | #N/A | #N/A | #N/A | Discolored Jaws | 0
Average Yield: |  |  | 90.89%

## EB016-EB216
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description |  | Comments
1486757 | EB216 | 1152 | 1192 | 1134 | 95.13% | 2023-04-18 | [15 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 10 |  | 2 | [8 empty] | 4 | 16 | 1.342% | 1192 | Stuck Blade | 16 | **Nut Setting Shim: .010**
[7 empty] | 1 | [10 empty] | 1 | 2 | 0.1678% | 1192 | Rough Actuation of Jaw | 2 | **Laser Welding Shim: .007**
[7 empty] | 2 |  | 1 | [9 empty] | 3 | 0.2517% | 1192 | 0.003 Jaw Gap Fail | 3
[7 empty] | 4 | [10 empty] | 1 | 5 | 0.4195% | 1192 | 0.006 Jaw Gap Fail | 5
[7 empty] | 3 | [11 empty] | 3 | 0.2517% | 1192 | Front Stops Not Touching | 3
[19 empty] | 0 | 0% | 1192 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1192 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1192 | Blade Does Not Pass Front Stop | 0
[7 empty] | 8 | [10 empty] | 5 | 13 | 1.091% | 1192 | Jaw Short | 13
[7 empty] | 2 | [11 empty] | 2 | 0.1678% | 1192 | Insulation Damage | 2
[7 empty] | 1 | [11 empty] | 1 | 0.08389% | 1192 | Incorrect Insulation Orientation | 1
[19 empty] | 0 | 0% | 1192 | Trigger Actuation | 0
[19 empty] | 0 | 0% | 1192 | Knob Collar Pressed incorrect | 0
[17 empty] | 1 |  | 1 | 0.08389% | 1192 | Partial Pack | 1
[19 empty] | 0 | 0% | 1192 | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 1192 | Damaged harness | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 1192 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1192 | Squeaky Blade | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 1192 | Continuity Fail | 0
[8 empty] | 4 | 1 | [9 empty] | 1 | 0.08389% | 1192 | High Jaw Force | 1
[8 empty] | 3 | [10 empty] | 0 | 0% | 1192 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 1192 | Return Spring Pop Out | 0
[19 empty] | 0 | 0% | 1192 | Fuse Switch | 0
[19 empty] | 0 | 0% | 1192 | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 1192 | Missing Component | 0 | BP Cap Separation: ea
[8 empty] | 4 | 2 | [9 empty] | 2 | 0.1678% | 1192 | Misassembled | 2 | Damaged Front Blade Pusher: 48ea
[8 empty] | 1 | [10 empty] | 0 | 0% | 1192 | Cosmetic Defect | 0 | 109ea Damaged Insulation
[19 empty] | 0 | 0% | 1192 | Scratched Logo | 0
[19 empty] | 0 | 0% | 1192 | Particulate Matter | 0
Shafts
[19 empty] | 0 | 0% | 1192 | Damaged Nut Theard | 0 | 11ea Incorrect Blade Stop Crimp ( No Jaw)
[19 empty] | 0 | 0% | 1192 | Damaged Insulation | 0 | 2ea Incorrect Hub Press
[7 empty] | 2 | [11 empty] | 2 | 0.1678% | 1192 | Exposed Wire | 2 | 14ea Damaged Pull Tube
[19 empty] | 0 | 0% | 1192 | Incorrect Weld | 0 | 4ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 1192 | Misassembled | 0
[19 empty] | 0 | 0% | 1192 | Damaged Blade | 0
[19 empty] | 0 | 0% | 1192 | Stuck Blade | 0
[7 empty] | 7 | [11 empty] | 7 | 0.5872% | 1192 | Blade Pusher Coming Out | 7
[6 empty] | Total | 40 | 15 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 11 | 58 | 4.866% | 1192
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description |  | Comments
1486758 | EB216 | 1920 | 2067 | 1855 | 89.74% | 2023-05-02 | [15 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 35 |  | 6 | [8 empty] | 28 | 69 | 3.338% | 2067 | Stuck Blade | 69 | **Nut Setting Shim: .010**
[7 empty] | 37 | [10 empty] | 7 | 44 | 2.129% | 2067 | Rough Actuation of Jaw | 44 | **Laser Welding Shim: .007**
[7 empty] | 31 |  | 1 | [8 empty] | 2 | 34 | 1.645% | 2067 | 0.003 Jaw Gap Fail | 34
[7 empty] | 2 | [10 empty] | 6 | 8 | 0.387% | 2067 | 0.006 Jaw Gap Fail | 8
[7 empty] | 4 | [11 empty] | 4 | 0.1935% | 2067 | Front Stops Not Touching | 4
[19 empty] | 0 | 0% | 2067 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2067 | Front and Middle Stops Not Touching | 0
[7 empty] | 1 |  | 1 | [9 empty] | 2 | 0.09676% | 2067 | Blade Does Not Pass Front Stop | 2
[7 empty] | 5 | [11 empty] | 5 | 0.2419% | 2067 | Jaw Short | 5
[7 empty] | 4 |  | 1 | [8 empty] | 3 | 8 | 0.387% | 2067 | Insulation Damage | 8
[7 empty] | 19 | [9 empty] | 17 |  | 36 | 1.742% | 2067 | Incorrect Insulation Orientation | 36
[19 empty] | 0 | 0% | 2067 | Trigger Actuation | 0
[7 empty] | 3 | [11 empty] | 3 | 0.1451% | 2067 | Knob Collar Pressed incorrect | 3
[17 empty] | 2 |  | 2 | 0.09676% | 2067 | Partial Pack | 2
[19 empty] | 0 | 0% | 2067 | Incomplete / Incorrect Weld | 0
[7 empty] | 1 |  | 2 | [9 empty] | 3 | 0.1451% | 2067 | Damaged harness | 3
[8 empty] | 7 | [10 empty] | 0 | 0% | 2067 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2067 | Squeaky Blade | 0
[8 empty] | 3 | [9 empty] | 5 | 5 | 0.2419% | 2067 | Continuity Fail | 5
[8 empty] | 8 | 1 | [9 empty] | 1 | 0.04838% | 2067 | High Jaw Force | 1
[8 empty] | 2 | [10 empty] | 0 | 0% | 2067 | Low Jaw Force | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2067 | Return Spring Pop Out | 0
[8 empty] | 1 | 1 | [9 empty] | 1 | 0.04838% | 2067 | Fuse Switch | 1
[8 empty] | 2 | [10 empty] | 0 | 0% | 2067 | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 2067 | Missing Component | 0 | BP Cap Separation: 10ea
[8 empty] | 6 | [10 empty] | 0 | 0% | 2067 | Misassembled | 0 | Damaged Front Blade Pusher: 192ea
[8 empty] | 2 | [10 empty] | 0 | 0% | 2067 | Cosmetic Defect | 0 | 86ea Damaged Insulation
[8 empty] | 1 | [9 empty] | 1 | 1 | 0.04838% | 2067 | Scratched Logo | 1
[19 empty] | 0 | 0% | 2067 | Particulate Matter | 0
[22 empty] | Shafts |  | 4ea Double Blade
[7 empty] | 2 | [11 empty] | 2 | 0.09676% | 2067 | Incorrect Jaw Crimp | 2 | 5ea Incorrect Blade Stop Crimp ( No Jaw)
[19 empty] | 0 | 0% | 2067 | Damaged Insulation | 0 | 1ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.04838% | 2067 | Exposed Wire | 1 | 3ea Misassembled (No Blade)
[19 empty] | 0 | 0% | 2067 | Incorrect Weld | 0 | 11ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2067 | Misassembled | 0 | 6ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04838% | 2067 | Damaged Component | 1 | 11ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2067 | Stuck Blade | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04838% | 2067 | Blade Pusher Coming Out | 1
[6 empty] | Total | 147 | 36 | 13 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 19 | 52 | 231 | 11.18% | 2067
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description |  | Comments
1486337 | EB216 | 1152 | 1249 | 1128 | 90.31% | 2023-05-09 | [15 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 37 |  | 2 | [8 empty] | 3 | 42 | 3.363% | 1249 | Stuck Blade | 42 | **Nut Setting Shim: .010**
[7 empty] | 10 |  | 1 | [9 empty] | 11 | 0.8807% | 1249 | Rough Actuation of Jaw | 11 | **Laser Welding Shim: .007**
[7 empty] | 12 |  | 1 | [8 empty] | 5 | 18 | 1.441% | 1249 | 0.003 Jaw Gap Fail | 18
[7 empty] | 4 | [11 empty] | 4 | 0.3203% | 1249 | 0.006 Jaw Gap Fail | 4
[7 empty] | 2 | [11 empty] | 2 | 0.1601% | 1249 | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 1249 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1249 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1249 | Blade Does Not Pass Front Stop | 0
[7 empty] | 8 |  | 2 | [8 empty] | 2 | 12 | 0.9608% | 1249 | Jaw Short | 12
[7 empty] | 4 | [11 empty] | 4 | 0.3203% | 1249 | Insulation Damage | 4
[7 empty] | 8 | [10 empty] | 4 | 12 | 0.9608% | 1249 | Incorrect Insulation Orientation | 12
[7 empty] | 1 | [11 empty] | 1 | 0.08006% | 1249 | Cosmetic Defect | 1 | Logo Missing Ink
[7 empty] | 1 | [11 empty] | 1 | 0.08006% | 1249 | Knob Collar Pressed incorrect | 1
[19 empty] | 0 | 0% | 1249 | Partial Pack | 0
[19 empty] | 0 | 0% | 1249 | Incomplete / Incorrect Weld | 0
[9 empty] | 2 | [9 empty] | 2 | 0.1601% | 1249 | Damaged harness | 2
[19 empty] | 0 | 0% | 1249 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1249 | Squeaky Blade | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 1249 | Continuity Fail | 0
[8 empty] | 7 | [10 empty] | 0 | 0% | 1249 | High Jaw Force | 0
[19 empty] | 0 | 0% | 1249 | Low Jaw Force | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 1249 | Return Spring Pop Out | 0
[19 empty] | 0 | 0% | 1249 | Fuse Switch | 0
[19 empty] | 0 | 0% | 1249 | Blade Lever Stop Pop Out | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 1249 | Missing Component | 0 | BP Cap Separation: ea
[8 empty] | 12 | 1 | [9 empty] | 1 | 0.08006% | 1249 | Misassembled | 1 | Damaged Front Blade Pusher: 88ea
[8 empty] | 5 | 1 | [9 empty] | 1 | 0.08006% | 1249 | Cosmetic Defect | 1 | Damaged Insulation: 128ea
[8 empty] | 1 | [10 empty] | 0 | 0% | 1249 | Scratched Logo | 0
[19 empty] | 0 | 0% | 1249 | Particulate Matter | 0
Shafts
[19 empty] | 0 | 0% | 1249 | Incorrect Jaw Crimp | 0
[7 empty] | 4 | [11 empty] | 4 | 0.3203% | 1249 | Insulation Damaged | 4 | 67ea Damaged Upper Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.1601% | 1249 | Exposed Wire | 2 | 6ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.08006% | 1249 | Incorrect Weld | 1 | 6ea Defective Pull Tube
[19 empty] | 0 | 0% | 1249 | Misassembled | 0 | 1ea Incorrect Blade Orientation
[7 empty] | 3 | [11 empty] | 3 | 0.2402% | 1249 | Defective Component | 3 | 1ea Blade Pusher Coming Out
[19 empty] | 0 | 0% | 1249 | Stuck Blade | 0
[19 empty] | 0 | 0% | 1249 | Blade Pusher Coming Out | 0
[6 empty] | Total | 97 | 31 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 14 | 121 | 9.688% | 1249
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description |  | Comments
1488010 | EB216 | 1152 | 1259 | 1104 | 87.69% | 2023-05-19 | [15 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 42 |  | 5 | [8 empty] | 20 | 67 | 5.322% | 1259 | Stuck Blade | 67 | **Nut Setting Shim: .010**
[7 empty] | 25 |  | 1 | [8 empty] | 2 | 28 | 2.224% | 1259 | Rough Actuation of Jaw | 28 | **Laser Welding Shim: .007**
[7 empty] | 10 |  | 1 | [9 empty] | 11 | 0.8737% | 1259 | 0.003 Jaw Gap Fail | 11
[7 empty] | 1 |  | 1 | [9 empty] | 2 | 0.1589% | 1259 | 0.006 Jaw Gap Fail | 2
[19 empty] | 0 | 0% | 1259 | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 1259 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1259 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1259 | Blade Does Not Pass Front Stop | 0
[7 empty] | 1 | [10 empty] | 6 | 7 | 0.556% | 1259 | Jaw Short | 7
[7 empty] | 2 | [11 empty] | 2 | 0.1589% | 1259 | Insulation Damage | 2
[7 empty] | 22 | [10 empty] | 2 | 24 | 1.906% | 1259 | Incorrect Insulation Orientation | 24
[19 empty] | 0 | 0% | 1259 | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 1259 | Knob Collar Pressed incorrect | 0
[17 empty] | 1 |  | 1 | 0.07943% | 1259 | Partial Pack | 1
[7 empty] | 1 | [11 empty] | 1 | 0.07943% | 1259 | Incomplete / Incorrect Weld | 1
[9 empty] | 4 | [9 empty] | 4 | 0.3177% | 1259 | Damaged harness | 4
[8 empty] | 2 | [10 empty] | 0 | 0% | 1259 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1259 | Squeaky Blade | 0
[18 empty] | 2 | 2 | 0.1589% | 1259 | Continuity Fail | 2
[8 empty] | 17 | 1 | [9 empty] | 1 | 0.07943% | 1259 | High Jaw Force | 1
[19 empty] | 0 | 0% | 1259 | Low Jaw Force | 0
[19 empty] | 0 | 0% | 1259 | Return Spring Pop Out | 0
[19 empty] | 0 | 0% | 1259 | Fuse Switch | 0
[19 empty] | 0 | 0% | 1259 | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 1259 | Missing Component | 0 | BP Cap Separation: 2ea
[8 empty] | 3 | 3 | [9 empty] | 3 | 0.2383% | 1259 | Misassembled | 3 | Damaged Front Blade Pusher: 53ea
[8 empty] | 5 | [10 empty] | 0 | 0% | 1259 | Cosmetic Defect | 0
[18 empty] | 1 | 1 | 0.07943% | 1259 | Scratched Logo | 1
[19 empty] | 0 | 0% | 1259 | Particulate Matter | 0
Shafts
[7 empty] | 1 | [11 empty] | 1 | 0.07943% | 1259 | Incorrect Jaw Crimp | 1
[19 empty] | 0 | 0% | 1259 | Insulation Damaged | 0 | 14ea Incorrect Blade Stop Crimp
[19 empty] | 0 | 0% | 1259 | Exposed Wire | 0 | 4ea Incorrect Hub Press
[19 empty] | 0 | 0% | 1259 | Incorrect/Missing Weld | 0 | 8ea Damaged Pull Tube
[19 empty] | 0 | 0% | 1259 | Misassembled | 0 | 33ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 1259 | Defective Component | 0
[19 empty] | 0 | 0% | 1259 | Stuck Blade | 0
[7 empty] | 2 | [11 empty] | 2 | 0.1589% | 1259 | Blade Pusher Coming Out | 2
[6 empty] | Total | 107 | 27 | 16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 33 | 157 | 12.47% | 1259
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | Description |  | Comments
1486338 | EB216 | 1152 | 1225 | 1122 | 91.59% | 2023-06-07 | [15 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[9 empty] | 1 | [8 empty] | 16 | 17 | 1.388% | 1225 | Stuck Blade | 17 | **Nut Setting Shim: .010**
[7 empty] | 15 | [11 empty] | 15 | 1.224% | 1225 | Rough Actuation of Jaw | 15 | **Laser Welding Shim: .007**
[7 empty] | 15 |  | 1 | [9 empty] | 16 | 1.306% | 1225 | 0.003 Jaw Gap Fail | 16
[7 empty] | 6 | [11 empty] | 6 | 0.4898% | 1225 | 0.006 Jaw Gap Fail | 6
[19 empty] | 0 | 0% | 1225 | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 1225 | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1225 | Front and Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1225 | Blade Does Not Pass Front Stop | 0
[19 empty] | 0 | 0% | 1225 | Jaw Short | 0
[7 empty] | 7 | [11 empty] | 7 | 0.5714% | 1225 | Insulation Damage | 7
[7 empty] | 11 | [11 empty] | 11 | 0.898% | 1225 | Incorrect Insulation Orientation | 11
[19 empty] | 0 | 0% | 1225 | Cosmetic Defect | 0
[7 empty] | 1 | [11 empty] | 1 | 0.08163% | 1225 | Knob Collar Pressed incorrect | 1
[17 empty] | 4 |  | 4 | 0.3265% | 1225 | Partial Pack | 4
[19 empty] | 0 | 0% | 1225 | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 1225 | Damaged harness | 0
[19 empty] | 0 | 0% | 1225 | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1225 | Squeaky Blade | 0
[8 empty] | 3 | 3 | [9 empty] | 3 | 0.2449% | 1225 | Continuity Fail | 3
[8 empty] | 36 | 2 | [9 empty] | 2 | 0.1633% | 1225 | High Jaw Force | 2
[9 empty] | 1 | [9 empty] | 1 | 0.08163% | 1225 | Low Jaw Force | 1
[19 empty] | 0 | 0% | 1225 | Return Spring Pop Out | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 1225 | Fuse Switch | 0
[19 empty] | 0 | 0% | 1225 | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 1225 | Missing Component | 0 | BP Cap Separation:
[8 empty] | 3 | [10 empty] | 0 | 0% | 1225 | Misassembled | 0 | Damaged Front Blade Pusher: 10ea
[8 empty] | 4 | [10 empty] | 0 | 0% | 1225 | Cosmetic Defect | 0
[17 empty] | 2 |  | 2 | 0.1633% | 1225 | Scratched Logo | 2
[19 empty] | 0 | 0% | 1225 | Particulate Matter | 0
Shafts
[19 empty] | 0 | 0% | 1225 | Incorrect Jaw Crimp | 0
[7 empty] | 4 | [11 empty] | 4 | 0.3265% | 1225 | Damaged Blade | 4 | 15ea Incorrect Blade Stop Crimp (No Jaw)
[7 empty] | 3 | [11 empty] | 3 | 0.2449% | 1225 | Exposed Wire | 3 | 11ea Incorrect Hub Press
[7 empty] | 5 | [11 empty] | 5 | 0.4082% | 1225 | Incorrect/Missing Weld | 5 | 14ea Damaged Pull Tube
[19 empty] | 0 | 0% | 1225 | Misassembled | 0 | 9ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 1225 | Defective Component | 0
[19 empty] | 0 | 0% | 1225 | Stuck Blade | 0
[7 empty] | 6 | [11 empty] | 6 | 0.4898% | 1225 | Blade Pusher Coming Out | 6
[6 empty] | Total | 73 | 47 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 16 | 103 | 8.408% | 1225

## EB017-EB217 Graphs
EB017/EB217 Energy Hand Device Yield
Scrap Breakdown (Last 3 S/O)
[5 empty] | Description | Count
[5 empty] | Stuck Blade | 416
[5 empty] | High Jaw Force | 66
[5 empty] | Rough Actuation of Jaw | 42
[5 empty] | 0.003 Jaw Gap Fail | 42
[5 empty] | Incorrect Insulation Orientation | 38
[5 empty] | Insulation Damage | 36
[5 empty] | Jaw Short | 10
[5 empty] | Continuity Fail | 6
[5 empty] | Front Stops Not Touching | 6
[5 empty] | Damaged Component | 3
[5 empty] | Low Jaw Force | 1
[5 empty] | Rough Knob Actuation | 1
[5 empty] | Misassembled | 0
[5 empty] | Fuse Switch | 0
[5 empty] | Blade Activation | 0
[5 empty] | Middle Stops Not Touching | 0
Shop Order's Yield |  |  |  |  | Rough Trigger | 0
Shop Order | Build QTY | Yield | Yield % | Date | Squeaky Blade | 0
1488011 | 1253 | 1123 | 89.62% | 2023-05-08 | Both Stops Not Touching | 0
1488946 | 2139 | 1839 | 85.97% | 2023-05-10 | 0.005 Jaw Gap Fail | 0
1491607 | 2260 | 1824 | 80.71% | 2023-05-23 | Damaged Jaw | 0
1491608 | 2174 | 1848 | 85% | 2023-06-02 | Damaged Handle | 0
1486763 | 1268 | 1125 | 88.72% | 2023-06-12 | Collapsed Front Stop | 0
1493293 | 2195 | 1822 | 83.01% | 2023-06-22 | Discolored Jaws | 0
Average Yield: |  |  | 85.51%

## EB017-EB217
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1484418 | EB217 | 1152 | 1213 | 1111 | 91.59% | 2023-03-30 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 34 |  | 1 | [8 empty] | 10 | 45 | 3.71% | 1213 |  | Stuck Blade | 45
[7 empty] | 3 | [10 empty] | 1 | 4 | 0.3298% | 1213 |  | Rough Actuation of Jaw | 4 | **Laser Welding Shim: .007 Shim**
[7 empty] | 3 |  | 2 | [8 empty] | 15 | 20 | 1.649% | 1213 |  | 0.003 Jaw Gap Fail | 20 | **Nut Setting Shim: .010 Shim**
[7 empty] | 2 | [10 empty] | 2 | 4 | 0.3298% | 1213 |  | 0.006 Jaw Gap Fail | 4
[7 empty] | 1 | [11 empty] | 1 | 0.08244% | 1213 |  | Front Stops Not Touching | 1
[19 empty] | 0 | 0% | 1213 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1213 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 1213 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 2 | [10 empty] | 1 | 3 | 0.2473% | 1213 |  | Jaw Short | 3
[7 empty] | 6 |  | 1 | [8 empty] | 1 | 8 | 0.6595% | 1213 |  | Insulation Damage | 8
[19 empty] | 0 | 0% | 1213 |  | Incorrect Insulation Orientation | 0
[19 empty] | 0 | 0% | 1213 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 1213 |  | Missing Harness
[19 empty] | 0 | 0% | 1213 |  | Damaged Jaw
[19 empty] | 0 | 0% |  |  | Cosmetic Defect
[9 empty] | 4 | [9 empty] | 4 | 0.3298% | 1213 |  | Damaged Harness | 4
[19 empty] | 0 | 0% | 1213 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1213 |  | Squeaky Blade | 0
[8 empty] | 2 | [9 empty] | 1 | 1 | 0.08244% | 1213 |  | Continuity Fail | 1
[8 empty] | 6 | [10 empty] | 0 | 0% | 1213 |  | High Jaw Force | 0
[9 empty] | 1 | [9 empty] | 1 | 0.08244% | 1213 |  | Low Jaw Force | 1
[8 empty] | 3 | [10 empty] | 0 | 0% | 1213 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[8 empty] | 1 | [10 empty] | 0 | 0% | 1213 |  | Fuse Switch | 0 | Damaged Front Blade Pusher : 20ea
[19 empty] | 0 | 0% | 1213 |  | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 1213 |  | Scratched Logo | 0
[8 empty] | 7 | 1 | [9 empty] | 1 | 0.08244% | 1213 |  | Misassembled | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 1213 |  | Contamination | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 1213 |  | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 1213 |  | Partial Pack | 0
[23 empty] | Shafts | 0
[19 empty] | 0 | 0% | 1213 |  | Insulation Damage | 0
[19 empty] | 0 | 0% | 1213 |  | Damaged Nut Thread | 1
[7 empty] | 1 | [11 empty] | 1 | 0.08244% | 1213 |  | Exposed Wire | 0 | 3ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 1213 |  | Incorrect Blade Stop Crimp | 1 | 8ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.08244% | 1213 |  | Incomplete / Incorrect Weld | 0 | 6ea Incorrect Hub Press
[19 empty] | 0 | 0% | 1213 |  | Contamination | 0 | 6ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 1213 |  | Damaged Component | 8
[7 empty] | 8 | [11 empty] | 8 | 0.6595% | 1213 |  | Blade Pusher Coming Out
[6 empty] | Total | 61 | 21 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 31 | 102 | 8.409% | 1213
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1484419 | EB217 | 1152 | 1328 | 1062 | 79.97% | 2023-04-04 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 127 |  | 9 | [8 empty] | 37 | 173 | 13.03% | 1328 |  | Stuck Blade | 173
[7 empty] | 2 | [10 empty] | 2 | 4 | 0.3012% | 1328 |  | Rough Actuation of Jaw | 4 | **Laser Welding Shim: .007 Shim**
[7 empty] | 19 |  | 9 | [8 empty] | 6 | 34 | 2.56% | 1328 |  | 0.003 Jaw Gap Fail | 34 | **Nut Setting Shim: .010 Shim**
[19 empty] | 0 | 0% | 1328 |  | 0.006 Jaw Gap Fail | 0
[19 empty] | 0 | 0% | 1328 |  | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 1328 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1328 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 1328 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 1 | [11 empty] | 1 | 0.0753% | 1328 |  | Jaw Short | 1
[7 empty] | 1 |  | 1 | [9 empty] | 2 | 0.1506% | 1328 |  | Insulation Damage | 2
[7 empty] | 1 | [11 empty] | 1 | 0.0753% | 1328 |  | Incorrect Insulation Orientation | 1
[19 empty] | 0 | 0% | 1328 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 1328 |  | Missing Harness
[19 empty] | 0 | 0% | 1328 |  | Damaged Jaw
[9 empty] | 1 | [9 empty] | 1 | 0.0753% |  |  | Cosmetic Defect
[9 empty] | 4 | [9 empty] | 4 | 0.3012% | 1328 |  | Damaged Harness | 4
[19 empty] | 0 | 0% | 1328 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1328 |  | Squeaky Blade | 0
[8 empty] | 1 | 2 | [8 empty] | 2 | 4 | 0.3012% | 1328 |  | Continuity Fail | 4
[8 empty] | 7 | 3 | [9 empty] | 3 | 0.2259% | 1328 |  | High Jaw Force | 3
[8 empty] | 1 | 3 | [9 empty] | 3 | 0.2259% | 1328 |  | Low Jaw Force | 3
[19 empty] | 0 | 0% | 1328 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[19 empty] | 0 | 0% | 1328 |  | Fuse Switch | 0 | Damaged Blade Pusher : 25ea
[19 empty] | 0 | 0% | 1328 |  | Blade Lever Stop Pop Out | 0
[18 empty] | 1 | 1 | 0.0753% | 1328 |  | Scratched Logo | 1
[8 empty] | 5 | 7 | [9 empty] | 7 | 0.5271% | 1328 |  | Misassembled | 7
[19 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 1328 |  | Contamination | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 1328 |  | Cosmetic Defect | 3
[17 empty] | 3 |  | 3 | 0.2259% | 1328 |  | Partial Pack | 0
[23 empty] | Shafts | 0
[19 empty] | 0 | 0% | 1328 |  | Insulation Damage | 0
[19 empty] | 0 | 0% | 1328 |  | Damaged Nut Thread | 9
[7 empty] | 9 | [11 empty] | 9 | 0.6777% | 1328 |  | Exposed Wire | 0 | 3ea Incorrect Blade Stop Crimp (No Jaws)
[19 empty] | 0 | 0% | 1328 |  | Incorrect Blade Stop Crimp | 0 | 15ea Damaged Pull Tube
[19 empty] | 0 | 0% | 1328 |  | Incomplete / Incorrect Weld | 8
[7 empty] | 8 | [11 empty] | 8 | 0.6024% | 1328 |  | Misassembled | 1
[7 empty] | 1 | [11 empty] | 1 | 0.0753% | 1328 |  | Stuck Blade | 7
[7 empty] | 7 | [11 empty] | 7 | 0.5271% | 1328 |  | Blade Pusher Coming Out
[6 empty] | Total | 176 | 15 | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 48 | 266 | 20.03% | 1328
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1484420 | EB217 | 1152 | 1337 | 1121 | 83.84% | 2023-04-20 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 25 |  | 2 | [8 empty] | 17 | 44 | 3.291% | 1337 |  | Stuck Blade | 44
[7 empty] | 5 | [10 empty] | 1 | 6 | 0.4488% | 1337 |  | Rough Actuation of Jaw | 6 | **Laser Welding Shim: .007 Shim**
[7 empty] | 7 | [10 empty] | 2 | 9 | 0.6731% | 1337 |  | 0.003 Jaw Gap Fail | 9 | **Nut Setting Shim: .010 Shim**
[19 empty] | 0 | 0% | 1337 |  | 0.006 Jaw Gap Fail | 0
[7 empty] | 2 | [11 empty] | 2 | 0.1496% | 1337 |  | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 1337 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1337 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 1337 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 3 | [10 empty] | 1 | 4 | 0.2992% | 1337 |  | Jaw Short | 4
[7 empty] | 8 | [10 empty] | 1 | 9 | 0.6731% | 1337 |  | Insulation Damage | 9
[19 empty] | 0 | 0% | 1337 |  | Incorrect Insulation Orientation | 0
[19 empty] | 0 | 0% | 1337 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 1337 |  | Partial Pack
[19 empty] | 0 | 0% | 1337 |  | Damaged Jaw
[19 empty] | 0 | 0% |  |  | Cosmetic Defect
[19 empty] | 0 | 0% | 1337 |  | Damaged Harness | 0
[8 empty] | 5 | 3 | [9 empty] | 3 | 0.2244% | 1337 |  | Rough Knob Actuation | 3
[19 empty] | 0 | 0% | 1337 |  | Squeaky Blade | 0
[8 empty] | 2 | 1 | [9 empty] | 1 | 0.07479% | 1337 |  | Continuity Fail | 1
[8 empty] | 14 | 1 | [9 empty] | 1 | 0.07479% | 1337 |  | High Jaw Force | 1
[8 empty] | 4 | [10 empty] | 0 | 0% | 1337 |  | Low Jaw Force | 0
[19 empty] | 0 | 0% | 1337 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[19 empty] | 0 | 0% | 1337 |  | Fuse Switch | 0 | Damaged Blade Pusher : 70ea
[19 empty] | 0 | 0% | 1337 |  | Blade Lever Stop Pop Out | 0 | Damaged Insulation: 99 ea
[18 empty] | 1 | 1 | 0.07479% | 1337 |  | Scratched Logo | 1
[8 empty] | 6 | [10 empty] | 0 | 0% | 1337 |  | Misassembled | 0
[19 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 1337 |  | Contamination | 0
[19 empty] | 0 | 0% | 1337 |  | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 1337 |  | Partial Pack | 0
[23 empty] | Shafts | 4
[7 empty] | 4 | [11 empty] | 4 | 0.2992% | 1337 |  | Insulation Damage | 3
[7 empty] | 3 | [11 empty] | 3 | 0.2244% | 1337 |  | Blade Does Not Pass Front Stop | 83
[7 empty] | 83 | [11 empty] | 83 | 6.208% | 1337 |  | Incorrect Blade Stop Crimp | 7 | 281ea Incorrect Blade Stop Crimp (No Jaws)
[7 empty] | 7 | [11 empty] | 7 | 0.5236% | 1337 |  | Incorrect Jaw Crimp | 1 | 18ea Incorrect Hub Press
[7 empty] | 1 | [11 empty] | 1 | 0.07479% | 1337 |  | Damaged Component | 0 | 6ea Missing Dimple Press ( No Jaws )
[19 empty] | 0 | 0% | 1337 |  | Misassembled | 3 | 5ea Damaged Upper Jaw
[7 empty] | 3 | [11 empty] | 3 | 0.2244% | 1337 |  | Contamination | 34 | 13ea Damaged Pull Tube
[7 empty] | 34 | [11 empty] | 34 | 2.543% | 1337 |  | Blade Pusher Coming Out |  | 2ea Damaged Cover Tube
[6 empty] | Total | 185 | 31 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 23 | 215 | 16.08% | 1337
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1486340 | EB217 | 1920 | 2148 | 1815 | 84.5% | 2023-04-26 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 112 |  | 12 |  | 2 | [6 empty] | 71 | 197 | 9.171% | 2148 |  | Stuck Blade | 197
[7 empty] | 35 |  | 1 | [8 empty] | 4 | 40 | 1.862% | 2148 |  | Rough Actuation of Jaw | 40 | **Laser Welding Shim: .007 Shim**
[7 empty] | 15 |  | 2 | [8 empty] | 3 | 20 | 0.9311% | 2148 |  | 0.003 Jaw Gap Fail | 20 | **Nut Setting Shim: .010 Shim**
[18 empty] | 8 | 8 | 0.3724% | 2148 |  | 0.006 Jaw Gap Fail | 8
[7 empty] | 1 | [11 empty] | 1 | 0.04655% | 2148 |  | Front Stops Not Touching | 1
[19 empty] | 0 | 0% | 2148 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2148 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 2148 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 3 | [10 empty] | 2 | 5 | 0.2328% | 2148 |  | Jaw Short | 5
[7 empty] | 4 |  | 4 | [8 empty] | 7 | 15 | 0.6983% | 2148 |  | Insulation Damage | 15
[7 empty] | 7 | [11 empty] | 7 | 0.3259% | 2148 |  | Incorrect Insulation Orientation | 7
[19 empty] | 0 | 0% | 2148 |  | Incomplete / Incorrect Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04655% | 2148 |  | Exposed Wire
[7 empty] | 17 | [11 empty] | 17 | 0.7914% | 2148 |  | Knob Collar Press incorrectly
[19 empty] | 0 | 0% |  |  | Cosmetic Defect
[11 empty] | 1 | [7 empty] | 1 | 0.04655% | 2148 |  | Damaged Harness | 1
[8 empty] | 4 | [10 empty] | 0 | 0% | 2148 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2148 |  | Squeaky Blade | 0
[8 empty] | 1 | [9 empty] | 1 | 1 | 0.04655% | 2148 |  | Continuity Fail | 1
[8 empty] | 29 | [10 empty] | 0 | 0% | 2148 |  | High Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2148 |  | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2148 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[19 empty] | 0 | 0% | 2148 |  | Fuse Switch | 0 | Damaged Blade Pusher : 61ea
[19 empty] | 0 | 0% | 2148 |  | Blade Lever Stop Pop Out | 0
[18 empty] | 4 | 4 | 0.1862% | 2148 |  | Scratched Logo | 4
[8 empty] | 17 | [10 empty] | 0 | 0% | 2148 |  | Misassembled | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 2148 |  | Contamination | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 2148 |  | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 2148 |  | Partial Pack | 0
[23 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2148 |  | Insulation Damage | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04655% | 2148 |  | Exposed Wire | 0 | 5ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2148 |  | Incorrect Blade Stop Crimp | 2 | 5ea Damaged Upper Jaw
[7 empty] | 2 | [11 empty] | 2 | 0.09311% | 2148 |  | Incorrect Jaw Crimp | 2 | 9ea Damaged Pull Tube
[7 empty] | 2 | [11 empty] | 2 | 0.09311% | 2148 |  | Damaged Component | 0
[19 empty] | 0 | 0% | 2148 |  | Misassembled | 0
[19 empty] | 0 | 0% | 2148 |  | Contamination | 9
[7 empty] | 9 | [11 empty] | 9 | 0.419% | 2148 |  | Blade Pusher Coming Out
[6 empty] | Total | 209 | 55 | 19 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 100 | 331 | 15.41% | 2148
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1488011 | EB217 | 1152 | 1253 | 1123 | 89.62% | 2023-05-08 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 23 |  | 1 | [8 empty] | 11 | 35 | 2.793% | 1253 |  | Stuck Blade | 35
[7 empty] | 21 | [10 empty] | 3 | 24 | 1.915% | 1253 |  | Rough Actuation of Jaw | 24 | **Laser Welding Shim: .007 Shim**
[7 empty] | 17 | [10 empty] | 2 | 19 | 1.516% | 1253 |  | 0.003 Jaw Gap Fail | 19 | **Nut Setting Shim: .010 Shim**
[7 empty] | 11 | [10 empty] | 2 | 13 | 1.038% | 1253 |  | 0.006 Jaw Gap Fail | 13
[7 empty] | 1 | [11 empty] | 1 | 0.07981% | 1253 |  | Front Stops Not Touching | 1
[19 empty] | 0 | 0% | 1253 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1253 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 1253 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 4 | [11 empty] | 4 | 0.3192% | 1253 |  | Jaw Short | 4
[7 empty] | 8 | [10 empty] | 2 | 10 | 0.7981% | 1253 |  | Insulation Damage | 10
[7 empty] | 2 | [11 empty] | 2 | 0.1596% | 1253 |  | Incorrect Insulation Orientation | 2
[19 empty] | 0 | 0% | 1253 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 1253 |  | Exposed Wire
[7 empty] | 9 | [11 empty] | 9 | 0.7183% | 1253 |  | Knob Collar Press incorrectly
[19 empty] | 0 | 0% |  |  | Cosmetic Defect
[9 empty] | 5 | [9 empty] | 5 | 0.399% | 1253 |  | Damaged Harness | 5
[8 empty] | 6 | [10 empty] | 0 | 0% | 1253 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1253 |  | Squeaky Blade | 0
[8 empty] | 1 | [9 empty] | 1 | 1 | 0.07981% | 1253 |  | Continuity Fail | 1
[8 empty] | 13 | [10 empty] | 0 | 0% | 1253 |  | High Jaw Force | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 1253 |  | Low Jaw Force | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 1253 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[19 empty] | 0 | 0% | 1253 |  | Fuse Switch | 0 | Damaged Blade Pusher : 23ea
[8 empty] | 1 | [10 empty] | 0 | 0% | 1253 |  | Blade Lever Stop Pop Out | 0
[18 empty] | 1 | 1 | 0.07981% | 1253 |  | Scratched Logo | 1
[8 empty] | 13 | 1 | [9 empty] | 1 | 0.07981% | 1253 |  | Misassembled | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 1253 |  | Contamination | 0
[8 empty] | 2 | [10 empty] | 0 | 0% | 1253 |  | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 1253 |  | Partial Pack | 0
[23 empty] | Shafts | 0
[19 empty] | 0 | 0% | 1253 |  | Insulation Damage | 0 | 3ea Incorrect Blade Stop Crimp
[19 empty] | 0 | 0% | 1253 |  | Exposed Wire | 0 | 2ea Incorrect Hub Press
[19 empty] | 0 | 0% | 1253 |  | Incorrect Blade Stop Crimp | 0 | 7ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 1253 |  | Incorrect Jaw Crimp | 3 | 7ea Damaged Pull Tube
[7 empty] | 3 | [11 empty] | 3 | 0.2394% | 1253 |  | Damaged Component | 0 | 5ea Rough Cover tube
[19 empty] | 0 | 0% | 1253 |  | Misassembled | 0 | 2ea Incorrect Blade Orientation
[19 empty] | 0 | 0% | 1253 |  | Contamination | 2
[7 empty] | 2 | [11 empty] | 2 | 0.1596% | 1253 |  | Blade Pusher Coming Out
[6 empty] | Total | 101 | 40 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 22 | 130 | 10.38% | 1253
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1488946 | EB217 | 1920 | 2139 | 1839 | 85.97% | 2023-05-10 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 78 |  | 1 | [8 empty] | 40 | 119 | 5.563% | 2139 |  | Stuck Blade | 119
[7 empty] | 26 | [10 empty] | 8 | 34 | 1.59% | 2139 |  | Rough Actuation of Jaw | 34 | **Laser Welding Shim: .007 Shim**
[7 empty] | 39 |  | 4 | [8 empty] | 2 | 45 | 2.104% | 2139 |  | 0.003 Jaw Gap Fail | 45 | **Nut Setting Shim: .010 Shim**
[7 empty] | 4 | [10 empty] | 3 | 7 | 0.3273% | 2139 |  | 0.006 Jaw Gap Fail | 7
[7 empty] | 1 |  | 1 | [9 empty] | 2 | 0.0935% | 2139 |  | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 2139 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2139 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 2139 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 7 | [10 empty] | 4 | 11 | 0.5143% | 2139 |  | Jaw Short | 11
[7 empty] | 2 |  | 2 | [9 empty] | 4 | 0.187% | 2139 |  | Insulation Damage | 4
[7 empty] | 17 |  | 2 | [8 empty] | 3 | 22 | 1.029% | 2139 |  | Incorrect Insulation Orientation | 22
[7 empty] | 1 | [11 empty] | 1 | 0.04675% | 2139 |  | Incomplete / Incorrect Weld | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04675% | 2139 |  | Blade Pusher Coming Out
[19 empty] | 0 | 0% | 2139 |  | Knob Collar Press incorrectly
[19 empty] | 0 | 0% |  |  | Cosmetic Defect
[7 empty] | 1 |  | 3 | [9 empty] | 4 | 0.187% | 2139 |  | Damaged Harness | 4
[8 empty] | 3 | [10 empty] | 0 | 0% | 2139 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2139 |  | Squeaky Blade | 0
[8 empty] | 1 | 2 | [8 empty] | 3 | 5 | 0.2338% | 2139 |  | Continuity Fail | 5
[8 empty] | 40 | 2 | [9 empty] | 2 | 0.0935% | 2139 |  | High Jaw Force | 2
[8 empty] | 2 | [10 empty] | 0 | 0% | 2139 |  | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2139 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[8 empty] | 2 | [10 empty] | 0 | 0% | 2139 |  | Fuse Switch | 0 | Damaged Blade Pusher : 32ea
[19 empty] | 0 | 0% | 2139 |  | Blade Lever Stop Pop Out | 0
[18 empty] | 1 | 1 | 0.04675% | 2139 |  | Scratched Logo | 1
[8 empty] | 9 | [10 empty] | 0 | 0% | 2139 |  | Misassembled | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 2139 |  | Contamination | 0
[8 empty] | 5 | [10 empty] | 0 | 0% | 2139 |  | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 2139 |  | Partial Pack | 0
[23 empty] | Shafts | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04675% | 2139 |  | Insulation Damage | 3 | 5ea Incorrect Blade Stop Crimp
[7 empty] | 3 | [11 empty] | 3 | 0.1403% | 2139 |  | Exposed Wire | 0 | 9ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2139 |  | Incorrect Blade Stop Crimp | 7 | 6ea Damaged Upper Jaw
[7 empty] | 7 | [11 empty] | 7 | 0.3273% | 2139 |  | Incorrect Jaw Crimp | 1 | 11ea Damaged Pull Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04675% | 2139 |  | Damaged Component | 3 | 3ea Damaged Cover Tube
[7 empty] | 3 | [11 empty] | 3 | 0.1403% | 2139 |  | Rear Blade Pusher Broken | 0
[19 empty] | 0 | 0% | 2139 |  | Contamination | 27
[7 empty] | 27 | [11 empty] | 27 | 1.262% | 2139 |  | Blade Pusher Coming Out
[6 empty] | Total | 219 | 63 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 64 | 300 | 14.03% | 2139
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1491607 | EB217 | 1920 | 2260 | 1824 | 80.71% | 2023-05-23 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 121 |  | 79 |  | 1 | [6 empty] | 42 | 243 | 10.75% | 2260 |  | Stuck Blade | 243
[7 empty] | 46 |  | 3 | [8 empty] | 4 | 53 | 2.345% | 2260 |  | Rough Actuation of Jaw | 53 | **Laser Welding Shim: .007 Shim**
[7 empty] | 33 |  | 5 | [8 empty] | 2 | 40 | 1.77% | 2260 |  | 0.003 Jaw Gap Fail | 40 | **Nut Setting Shim: .010 Shim**
[7 empty] | 5 |  |  |  | 1 | [7 empty] | 6 | 0.2655% | 2260 |  | 0.006 Jaw Gap Fail | 6
[7 empty] | 4 | [11 empty] | 4 | 0.177% | 2260 |  | Front Stops Not Touching | 4
[19 empty] | 0 | 0% | 2260 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2260 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 2260 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 4 |  | 4 | [7 empty] | 1 | 3 | 12 | 0.531% | 2260 |  | Jaw Short | 12
[7 empty] | 19 |  | 8 |  | 2 | [6 empty] | 6 | 35 | 1.549% | 2260 |  | Insulation Damage | 35
[7 empty] | 2 | [11 empty] | 2 | 0.0885% | 2260 |  | Incorrect Insulation Orientation | 2
[19 empty] | 0 | 0% | 2260 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 2260 |  | Blade Pusher Coming Out
[7 empty] | 3 | [11 empty] | 3 | 0.1327% | 2260 |  | Knob Collar Press incorrectly
[17 empty] | 1 |  | 1 | 0.04425% |  |  | Cosmetic Defect
[11 empty] | 21 | [7 empty] | 21 | 0.9292% | 2260 |  | Damaged Harness | 21
[10 empty] | 1 | [8 empty] | 0 | 0% | 2260 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2260 |  | Squeaky Blade | 0
[8 empty] | 8 | 1 | 1 | 2 | [6 empty] | 1 | 4 | 0.177% | 2260 |  | Continuity Fail | 4
[8 empty] | 26 |  | 1 | [8 empty] | 0 | 0% | 2260 |  | High Jaw Force | 0
[8 empty] | 4 | [10 empty] | 0 | 0% | 2260 |  | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2260 |  | Return Spring Pop Out | 0 | BP Cap Seperation : 7ea
[8 empty] | 2 |  | 2 | 2 | [7 empty] | 2 | 0.0885% | 2260 |  | Fuse Switch | 2 | Damaged Blade Pusher : 65ea
[19 empty] | 0 | 0% | 2260 |  | Blade Lever Stop Pop Out | 0
[18 empty] | 1 | 1 | 0.04425% | 2260 |  | Scratched Logo | 1
[8 empty] | 23 |  | 2 | 2 | [7 empty] | 2 | 0.0885% | 2260 |  | Misassembled | 2
[19 empty] | 0 | 0% | Build QTY |  | Missing Component | 1
[17 empty] | 1 |  | 1 | 0.04425% | 2260 |  | Contamination | 1
[8 empty] | 10 | [8 empty] | 1 |  | 1 | 0.04425% | 2260 |  | Particulate Matter | 2
[17 empty] | 2 |  | 2 | 0.0885% | 2260 |  | Partial Pack | 0
[23 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2260 |  | Over Tripped | 2 | 3ea Incorrect Blade Stop Crimp
[7 empty] | 2 | [11 empty] | 2 | 0.0885% | 2260 |  | Exposed Wire | 0 | 29ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2260 |  | Incorrect Blade Stop Crimp | 0 | 10ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2260 |  | Incorrect Jaw Crimp | 0 | 2ea Damaged Pull Tube
[19 empty] | 0 | 0% | 2260 |  | Damaged Component | 1 | 7ea Damaged Cover Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04425% | 2260 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 2260 |  | Stuck Blade | 0
[19 empty] | 0 | 0% | 2260 |  | Blade Pusher Coming Out
[6 empty] | Total | 240 | 73 | 100 | 7 | 31 | 0 | 0 | 0 | 0 | 0 | 6 | 59 | 436 | 19.29% | 2260
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1491608 | EB217 | 1920 | 2174 | 1848 | 85% | 2023-06-02 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 114 |  | 3 | [8 empty] | 16 | 133 | 6.118% | 2174 |  | Stuck Blade | 133
[7 empty] | 22 |  | 3 | [8 empty] | 1 | 26 | 1.196% | 2174 |  | Rough Actuation of Jaw | 26 | **Laser Welding Shim: .007 Shim**
[7 empty] | 13 | [10 empty] | 1 | 14 | 0.644% | 2174 |  | 0.003 Jaw Gap Fail | 14 | **Nut Setting Shim: .010 Shim**
[7 empty] | 1 | [10 empty] | 1 | 2 | 0.092% | 2174 |  | 0.006 Jaw Gap Fail | 2
[19 empty] | 0 | 0% | 2174 |  | Front Stops Not Touching | 0
[19 empty] | 0 | 0% | 2174 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2174 |  | Knob Collar Press Incorrectly | 0
[7 empty] | 1 | [11 empty] | 1 | 0.046% | 2174 |  | Blade Does Not Pass Front Stop | 1
[18 empty] | 4 | 4 | 0.184% | 2174 |  | Jaw Short | 4
[7 empty] | 5 |  | 4 | [9 empty] | 9 | 0.414% | 2174 |  | Insulation Damage | 9
[7 empty] | 9 |  | 2 | [8 empty] | 4 | 15 | 0.69% | 2174 |  | Incorrect Insulation Orientation | 15
[19 empty] | 0 | 0% | 2174 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 2174 |  | Blade Pusher Coming Out
[7 empty] | 7 | [11 empty] | 7 | 0.322% | 2174 |  | Knob Collar Press incorrectly
[19 empty] | 0 | 0% |  |  | Missing Harness
[9 empty] | 11 | [9 empty] | 11 | 0.506% | 2174 |  | Damaged Harness | 11
[8 empty] | 3 | [10 empty] | 0 | 0% | 2174 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 2174 |  | Squeaky Blade | 0
[9 empty] | 1 | [9 empty] | 1 | 0.046% | 2174 |  | Continuity Fail | 1
[8 empty] | 200 | 12 | [9 empty] | 12 | 0.552% | 2174 |  | High Jaw Force | 12
[19 empty] | 0 | 0% | 2174 |  | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2174 |  | Return Spring Pop Out | 0 | BP Cap Seperation : ea
[8 empty] | 1 | [10 empty] | 0 | 0% | 2174 |  | Fuse Switch | 0 | Damaged Blade Pusher : 42ea
[8 empty] | 1 | [10 empty] | 0 | 0% | 2174 |  | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 2174 |  | Scratched Logo | 0
[8 empty] | 6 | [10 empty] | 0 | 0% | 2174 |  | Misassembled | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 2174 |  | Contamination | 2
[8 empty] | 9 | 1 | [8 empty] | 1 | 2 | 0.092% | 2174 |  | Cosmetic Defect | 2
[17 empty] | 2 |  | 2 | 0.092% | 2174 |  | Partial Pack | 0
[23 empty] | Shafts | 0
[19 empty] | 0 | 0% | 2174 |  | Over Tripped | 1 | 6ea Incorrect Blade Stop Crimp
[7 empty] | 1 | [11 empty] | 1 | 0.046% | 2174 |  | Exposed Wire | 0 | 15ea Incorrect Hub Press
[19 empty] | 0 | 0% | 2174 |  | Incorrect Blade Stop Crimp | 0 | 12ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2174 |  | Incorrect Jaw Crimp | 6 | 7ea Damaged Pull Tube
[7 empty] | 6 | [11 empty] | 6 | 0.276% | 2174 |  | Damaged Component | 0 | 8ea Damaged Cover Tube
[19 empty] | 0 | 0% | 2174 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 2174 |  | Stuck Blade | 75
[7 empty] | 75 | [11 empty] | 75 | 3.45% | 2174 |  | Blade Pusher Coming Out
[6 empty] | Total | 254 | 221 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 28 | 321 | 14.77% | 2174
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1486763 | EB217 | 1152 | 1268 | 1125 | 88.72% | 2023-06-12 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 61 | [10 empty] | 11 | 72 | 5.678% | 1268 |  | Stuck Blade | 72
[7 empty] | 3 | [11 empty] | 3 | 0.2366% | 1268 |  | Rough Actuation of Jaw | 3 | **Laser Welding Shim: .007 Shim**
[7 empty] | 4 |  | 1 | [9 empty] | 5 | 0.3943% | 1268 |  | 0.003 Jaw Gap Fail | 5 | **Nut Setting Shim: .010 Shim**
[7 empty] | 3 | [11 empty] | 3 | 0.2366% | 1268 |  | 0.006 Jaw Gap Fail | 3
[7 empty] | 2 | [11 empty] | 2 | 0.1577% | 1268 |  | Front Stops Not Touching | 2
[19 empty] | 0 | 0% | 1268 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 1268 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 1268 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 1 | [11 empty] | 1 | 0.07886% | 1268 |  | Jaw Short | 1
[7 empty] | 5 |  | 5 | [8 empty] | 5 | 15 | 1.183% | 1268 |  | Insulation Damage | 15
[7 empty] | 1 | [10 empty] | 1 | 2 | 0.1577% | 1268 |  | Incorrect Insulation Orientation | 2
[19 empty] | 0 | 0% | 1268 |  | Incomplete / Incorrect Weld | 0
[19 empty] | 0 | 0% | 1268 |  | Blade Pusher Coming Out
[19 empty] | 0 | 0% | 1268 |  | Knob Collar Press incorrectly
[19 empty] | 0 | 0% |  |  | Missing Harness
[9 empty] | 1 | [9 empty] | 1 | 0.07886% | 1268 |  | Damaged Harness | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 1268 |  | Rough Knob Actuation | 0
[19 empty] | 0 | 0% | 1268 |  | Squeaky Blade | 0
[18 empty] | 1 | 1 | 0.07886% | 1268 |  | Continuity Fail | 1
[8 empty] | 29 | 1 | [9 empty] | 1 | 0.07886% | 1268 |  | High Jaw Force | 1
[8 empty] | 1 | 1 | [9 empty] | 1 | 0.07886% | 1268 |  | Low Jaw Force | 1
[8 empty] | 1 | [10 empty] | 0 | 0% | 1268 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[19 empty] | 0 | 0% | 1268 |  | Fuse Switch | 0 | Damaged Blade Pusher :
[19 empty] | 0 | 0% | 1268 |  | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 1268 |  | Scratched Logo | 0
[8 empty] | 3 | [10 empty] | 0 | 0% | 1268 |  | Misassembled | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 1268 |  | Contamination | 0
[8 empty] | 1 | [10 empty] | 0 | 0% | 1268 |  | Cosmetic Defect | 0
[19 empty] | 0 | 0% | 1268 |  | Partial Pack | 0
[23 empty] | Shafts | 0
[19 empty] | 0 | 0% | 1268 |  | Over Tripped | 4 | 6ea Incorrect Blade Stop Crimp (no Jaw)
[7 empty] | 4 | [11 empty] | 4 | 0.3155% | 1268 |  | Exposed Wire | 0 | 10ea Incorrect Hub Press
[19 empty] | 0 | 0% | 1268 |  | Incorrect Blade Stop Crimp | 0 | 19ea Damaged Pull Tube
[19 empty] | 0 | 0% | 1268 |  | Incorrect Jaw Crimp | 3
[7 empty] | 3 | [11 empty] | 3 | 0.2366% | 1268 |  | Damaged Component | 3 | 8ea Damaged Upper Jaw
[7 empty] | 3 | [11 empty] | 3 | 0.2366% | 1268 |  | Incomplete / Incorrect Weld | 1
[7 empty] | 1 | [11 empty] | 1 | 0.07886% | 1268 |  | Stuck Blade | 25
[7 empty] | 25 | [11 empty] | 25 | 1.972% | 1268 |  | Blade Pusher Coming Out
[6 empty] | Total | 116 | 37 | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 143 | 11.28% | 1268
Shop Order | Model | S/O Qty | Build QTY | Yield | Yield % | Date | Production Initial build Inspection (Scrapped Shafts) | Production Initial build rework (Same Shaft) | 1st Rework (New Shafts) | 1st Rework (Same Shafts) | 2nd Rework (New Shafts) | 2nd Rework (Same Shafts) | 3rd Rework (New Shafts) | 3rd Rework (Same Shafts) | 4th Rework (New Shafts) | 4th Rework Same Shafts) | Packaging | Quality Control Inspection | Total | Scrap % | RSL Code | RSL Code | Description |  | Comments
1493293 | EB217 | 1920 | 2195 | 1822 | 83.01% | 2023-06-22 | [16 empty] | Top Level Non-Conformance |  | **Production Shop Order**
[7 empty] | 51 |  | 9 |  | 7 | [6 empty] | 44 | 111 | 5.057% | 2195 |  | Stuck Blade | 111
[7 empty] | 10 |  | 3 | [9 empty] | 13 | 0.5923% | 2195 |  | Rough Actuation of Jaw | 13 | **Laser Welding Shim: .007 Shim**
[7 empty] | 19 |  | 1 | [8 empty] | 3 | 23 | 1.048% | 2195 |  | 0.003 Jaw Gap Fail | 23 | **Nut Setting Shim: .010 Shim**
[7 empty] | 6 | [11 empty] | 6 | 0.2733% | 2195 |  | 0.006 Jaw Gap Fail | 6
[7 empty] | 4 | [11 empty] | 4 | 0.1822% | 2195 |  | Front Stops Not Touching | 4
[19 empty] | 0 | 0% | 2195 |  | Middle Stops Not Touching | 0
[19 empty] | 0 | 0% | 2195 |  | Knob Collar Press Incorrectly | 0
[19 empty] | 0 | 0% | 2195 |  | Blade Does Not Pass Front Stop | 0
[7 empty] | 4 |  | 1 | [9 empty] | 5 | 0.2278% | 2195 |  | Jaw Short | 5
[7 empty] | 2 |  | 9 | [8 empty] | 1 | 12 | 0.5467% | 2195 |  | Insulation Damage | 12
[7 empty] | 19 |  | 1 | [8 empty] | 1 | 21 | 0.9567% | 2195 |  | Incorrect Insulation Orientation | 21
[19 empty] | 0 | 0% | 2195 |  | Incomplete / Incorrect Weld | 0
[7 empty] | 1 | [11 empty] | 1 | 0.04556% | 2195 |  | Blade Pusher Coming Out
[7 empty] | 10 | [11 empty] | 10 | 0.4556% | 2195 |  | Knob Collar Press incorrectly
[19 empty] | 0 | 0% |  |  | Missing Harness
[11 empty] | 16 | [7 empty] | 16 | 0.7289% | 2195 |  | Damaged Harness | 16
[8 empty] | 12 | 1 | [9 empty] | 1 | 0.04556% | 2195 |  | Rough Knob Actuation | 1
[19 empty] | 0 | 0% | 2195 |  | Squeaky Blade | 0
[8 empty] | 4 | 2 | [8 empty] | 2 | 4 | 0.1822% | 2195 |  | Continuity Fail | 4
[8 empty] | 247 | 30 | 49 | 23 | [7 empty] | 53 | 2.415% | 2195 |  | High Jaw Force | 53
[19 empty] | 0 | 0% | 2195 |  | Low Jaw Force | 0
[19 empty] | 0 | 0% | 2195 |  | Return Spring Pop Out | 0 | BP Cap Seperation :
[19 empty] | 0 | 0% | 2195 |  | Fuse Switch | 0 | Damaged Blade Pusher : 56ea
[10 empty] | 2 | [8 empty] | 0 | 0% | 2195 |  | Blade Lever Stop Pop Out | 0
[19 empty] | 0 | 0% | 2195 |  | Scratched Logo | 0
[8 empty] | 21 |  | 1 | [8 empty] | 0 | 0% | 2195 |  | Misassembled | 0
[19 empty] | 0 | 0% | Build QTY |  | Missing Component | 0
[19 empty] | 0 | 0% | 2195 |  | Contamination | 0
[8 empty] | 1 |  | 2 | [8 empty] | 0 | 0% | 2195 |  | Cosmetic Defect | 0
[8 empty] | 2 |  | 1 | [8 empty] | 0 | 0% | 2195 |  | Particulate Matter | 0
[23 empty] | Shafts | 4
[7 empty] | 4 | [11 empty] | 4 | 0.1822% | 2195 |  | Rough Blade Pusher | 1 | 24ea Incorrect Blade Stop Crimp (no Jaw)
[7 empty] | 1 | [11 empty] | 1 | 0.04556% | 2195 |  | Exposed Wire | 7 | 2ea Incorrect Hub Press
[7 empty] | 7 | [11 empty] | 7 | 0.3189% | 2195 |  | Insulation Damage | 0 | 6ea Damaged Pull Tube
[19 empty] | 0 | 0% | 2195 |  | Incorrect Jaw Crimp | 1 | 2ea Damaged Cover Tube
[7 empty] | 1 | [11 empty] | 1 | 0.04556% | 2195 |  | Damaged Component | 0 | 13ea Damaged Upper Jaw
[19 empty] | 0 | 0% | 2195 |  | Incomplete / Incorrect Weld | 1
[7 empty] | 1 | [11 empty] | 1 | 0.04556% | 2195 |  | Contamination | 78
[7 empty] | 78 | [11 empty] | 78 | 3.554% | 2195 |  | Blade Pusher Coming Out
[6 empty] | Total | 218 | 287 | 57 | 55 | 46 | 0 | 0 | 0 | 0 | 0 | 0 | 51 | 372 | 16.95% | 2195

## Sheet1 (hidden)
|  | Model |  |  | General Information
Shop Order
Shop Order Qty.
Build Qty.
Top-Level Yield (%)
| RSL Code | Non Conformance | Initial Build (Scrap) | Initial Build (Reassembly) | 1st Reassembly (Scrap) | 1st Reassembly (Reassembly) | 2nd Reassembly (Scrap) | 2nd Reassembly (Reassembly) | 3rd Reassembly (Scrap) | 3rd Reassembly (Reassembly) | QC | Packaging | Total Scrap | Scrap % | Comments
TOP-LEVEL SCRAP |  | Stuck Blade | [10 empty] | 0 | #DIV/0!
|  | Blade Actiavtion | [10 empty] | 0 | #DIV/0!
|  | Rough Jaw Actuation | [10 empty] | 0 | #DIV/0!
|  | Low Jaw Gap | [10 empty] | 0 | #DIV/0!
|  | Low Jaw Gap | [10 empty] | 0 | #DIV/0!
|  | High Jaw Gap | [10 empty] | 0 | #DIV/0!
|  | Front Stops Not Touching | [10 empty] | 0 | #DIV/0!
|  | Middle Stops Not Touching | [10 empty] | 0 | #DIV/0!
|  | Both Stops Not Touching | [10 empty] | 0 | #DIV/0!
|  | Blade Doesn't Pass Front Stop | [10 empty] | 0 | #DIV/0!
|  | Jaw Short | [10 empty] | 0 | #DIV/0!
|  | Insulation Damage | [10 empty] | 0 | #DIV/0!
|  | Insulation Orientation | [10 empty] | 0 | #DIV/0!
|  | Damaged Harness | [10 empty] | 0 | #DIV/0!
TOP-LEVEL REASSEMBLY |  | Rough Knob Rotation | [10 empty] | 0 | #DIV/0!
|  | Continuity Fail | [10 empty] | 0 | #DIV/0!
|  | High Jaw Force | [10 empty] | 0 | #DIV/0!
|  | Low Jaw Force | [10 empty] | 0 | #DIV/0!
|  | Return Spring Pop Out | [10 empty] | 0 | #DIV/0!
|  | Fuse Switch | [10 empty] | 0 | #DIV/0!
|  | Blade Lever Stop Pop Out | [10 empty] | 0 | #DIV/0!
|  | Open Handle | [10 empty] | 0 | #DIV/0!
|  | Missassembled | [10 empty] | 0 | #DIV/0!
|  | Cosmetic Defect | [10 empty] | 0 | #DIV/0!
SUB-ASSEMBLY SCRAP |  | Defective Component | [10 empty] | 0 | #DIV/0!
|  | Defective Weld | [10 empty] | 0 | #DIV/0!
|  | Missassembled | [10 empty] | 0 | #DIV/0!
|  | Contamination | [10 empty] | 0 | #DIV/0!
|  | Totals | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0!